K935.11 Super resolution microscope – Request for Information

SRM microscope for use in microbiology BSL-3 laboratory

The new Life Science Building at Gaustad in Oslo, Norway (scheduled for completion in 2027) will house state-of-the-art infrastructure, equipment, and expertise for advanced, cutting-edge research and innovation.

We would like to understand what suppliers can offer to meet our needs, as outlined below.
We would also appreciate a preliminary proposal for a recommended setup based on the specifications provided, along with an estimated cost for budgeting purposes. Please include a suitable computer system in the cost estimate, as well as a comprehensive software package for image analysis.

Due to time constraints, we would like to arrange a meeting to discuss the proposed solution as soon as possible after receiving the written proposal and budget estimate.

Needs: What do we want to see/study/inspect?

Many bacterial structures (protein clusters, cell wall synthesis sites, cell division machinery) are in the order of approx 10–100 nm. We need 60/63x and 100x objectives, suitable for looking at microbes and host cells in SRM methods. One objective for overview.

| Cell type | Cell size | Comment |
| --- | --- | --- |
| Mycobacterium tuberculosis (Mtb) | 2–4 µm length, 0.2–0.5 µm width | Thick cell wall, cell membrane with lipids mycolic acids/complex cell envelope, drug-resistant Slow growth: 18 hrs per generation |
| Escherichia coli | 1–2 µm length, 0.5–1.0 µm width | Gramnegative, thin cell wall, DNA compaction in nucleoids |
| Neisseria meningitidis | diplococcus 0.6-1.0 µm diameter | Gramnegative, pilus retraction, capsule, cell-cell interactions Fast growth: 20 min per generation |
| Candida species | 4-6 µm length, 3–5 µm width | Yeast |
| Macrophages with intracellular Mtb or other mycobacteria | 10-30 µm diameter | Intracellular Co-localization |
| Epithelial host cells with adherence/invasion | 10-40 µm diameter | adherence/invasion with fluorescence, Co-localization |
| Host cells infected by viruses; models for viral infections | 10-40 µm diameter |  |

Functional needs: Use SRM to:

- Resolve protein clusters and nano-domains in membranes.

- Visualize DNA in cell division and DNA compaction in nucleoids

- Map the bacterial cytoskeleton (e.g. MreB, FtsZ) at high precision.

- Visualize cell wall synthesis patches, secretion systems, flagellar motors, etc.

- Follow dynamic processes (cell division, chromosome organization, phagosomes) in live bacteria and host cells.

At a resolution down to ∼100 nm, in a manner which is fast live-cell imaging, as well as stained fixed cells. Please state fps for full SRM images, as well as maximum fps achievable (as complete SRM images and at what camera resolution).

Resolution ∼10-20nm on fixed samples. (please state what the smallest resolution is with the 10x NA SRM method is).

Must have:

SRM = 10x NA-value.

SRM (2X NA-value) for both live view and for stained fixed cells

At least four laser lines for fluorescence and SRM, approx. 405, 488, 561 and 640.

Optional:

- TIRF to study processes near cell wall/cell membrane.

Please state if it is possible to include TIRF in the same microscope with the two SRM methods as well. If so, please add an additional rough cost estimate that includes the TIRF system, for budget purposes.

- Is it possible to have phase contrast as well on the microscope? (Ability to microscope bacteria/cells without fluorescence labelling).

General information:

The equipment will be installed in a BSL-3 laboratory, outside a hood or safety cabinet, and must therefore be an incubator with controlled temperature, humidity, and CO₂ levels. A cage-type design is preferred, and the system should ideally include HEPA filtration on both the air inlet and outlet.

General description and examples:

Live cells: With SIM: Detection of live bacterial cells, in stationary phase or cell division, change of cellular shape and character/contrast, labelled with fluorescens either genetically or stained, 1-3 colors. Molecular co-localization with proteins labelled with different fluorescence tags during cellular processes and infection models.

Fast versus slow movements

Interaction with host cells

- Adherence / invasion of host epithelial cells

- Macrophages/host cells with intracellular Mtb or other mycobacteria

- Studies of cellular responses to viral infections

Demanding live view:

We aim to study fast processes like Neisserial pilus retraction and pilus-mediated motility and cellular invasion with SIM microscopy. For this purpose, it’s preferable to have a fast frame rate and still be able to operate with a reasonable resolution. In this context, we would like information on the achievable frame rate (fps) for the final super-resolution (SRM) images.

Stained fixed cells: SRM and regular FL imaging: labelled with fluorescens either genetically or stained, 1-3 colors.

Example: Microscopy of DNA in fixed cells

The DNA compaction phenotype of strains will be investigated with chromosomally expressed HU-mCherry or using conventional DNA staining with Hoechst 33258.

Fluorophores commonly used for labeling bacteria in our lab:

- Genetically encoded fluorescent proteins: GFP, mNeonGreen, mCherry, mEos3.2, PAmCherry.

- Organic dyes and self-labeling tags:

- Self-labeling tags: HaloTag, SNAP-tag, CLIP-tag fused to proteins.

- Small-molecule fluorophores: Alexa Fluor 647, CF dyes, Atto dyes, etc.

Problem: Permeability in bacteria can be tricky.

- Direct labeling of cell structures

- Cell wall: Fluorescent D-amino acids to label peptidoglycan synthesis.

- Membranes: Lipophilic dyes (e.g. FM dyes, DiI, Nile Red analogs).

- DNA: DAPI, Hoechst, or specific DNA-binding fluorophores.

- Surface proteins: Fluorescently labeled antibodies (more for fixed cells).

Question: please state if there are any limitations with any of the SRM methods regarding fluorescence labeling.

Please send your response by email to tom.ask@nosyko.no

Kindly submit your response no later than Friday, 29 May at 12:00 (Oslo time).

We look forward to hearing from you.