## 1. Guidance for use
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Project 1105601 UiO Livsvitenskap brukerutstyr
K935.10 Konfokal mikroskop
Tenderer:
HOW TO FILL OUT THE DOCUMENT
Please note that there are several tabs to be filled out.
Columns and rows are not to be deleted.
The tenderer shall only fill out fields with a white background. Fields of other colors are not to be used or altered by the Tenderer.
The tenderer shall answer YES or NO to the question on whether or not a requirement is fulfilled. All rows must be answered.
The tenderer shall answer the defined requirements. The answer shall have sufficient information for the contractor to carry out a full assessment of the tender without the need to open attachments. If the tenderer needs to provide additional or more detailed information than this document allows for, it is allowed to refer to attachments. In case of reference to attachments, the tenderer shall provide the corresponding attachment nr and relevant pages.
Minimum requirements (A-criteria) are not related to an award criterion. Lack of fulfillment of minimum requirement can lead to rejecton of the tender, cf Tilbudsinvitasjon (Tender invitation).
B-criteria are related to an award criterion. The tenderers answer to B-criteria are the basis of the contractors evaluation and assessment. B-criteria are defined as B1, B2 or B3 criteria in the corresponding column for each criterion. The definitions indicate the assigned weight of the award criterion in the tender evaluation. Award criteria indicated as B3 are assigned most weight, while the B1 criteria are assigned the least weight.
Information criteria (I-criteria) are solely for informative purposes, and are not part of the evaluation.
| Valuta: | Lisens:
| NOK | Systemlisens
| SEK | Brukerlisens
| DKK | Utstyrslisens
| EUR | Analyselisens
| USD | Annet
GBP

## 2. Point Scanning Confocal
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Navn på artikkel | Art. 025.02.001:3 Konfokalmikroskop, Point Scanning Confocal, opprett
Criterion nr | Description of criterion | Criterion type | Supplementary information/Additional information | Fulfillment of criterion Yes/N0 | Supplier response
1 | QUALITY AND FUNCTIONALITY CRITERIA
1.1 | The equipment shall be CE marked in accordance with current directives and fulfill all relevant Norwegian laws, standards, requirements and norms for the equipment offered. Declaration of conformity for CE marking (Declaration of conformity) shall be attached to the offer. | A
1.2 | Intended use: The system shall be a state-of-the-art multimodal, fully automated, upright, point scanning confocal microscope with high optical resolution, sensitivity, and imaging speed. It shall enable detailed visualization and quantification of subcellular structures, as well as imaging of small animals (e.g., zebrafish, C. elegans) and tissues. Fast scanning is required for non-invasive live 3D/4D imaging, supported by accurate data acquisition and analysis across a wide range of biomedical applications. | A | The microscope shall provide modular flexibility, intuitive software integration, and future upgrade options to accommodate evolving research needs. Intended samples: Live cells, organoids and embryos, thick tissue sections, live tissue slices, fixed specimens, cleared samples, as well as opaque samples in material research.
1.3 | State the estimated delivery time from when the order is received from the customer until the equipment is delivered. | A
1.4 | A complete user manual and safety data sheet in a Scandinavian language or English shall accompany the delivery, preferably in electronic format. | A
1.5 | The Tenderer shall, in addition, submit a detailed breakdown of all components included in the main offer, presented in separate line items including pricing, as a separate attachment. This is intended to provide an overview of all components included in the main offer. It does not replace Appendix 3.0 – Price Form. | A
Microscope Frame and Sample Environment
1.6 | The instrument shall be a fully motorized upright point-scanning confocal microscope with a motorized Z-drive and nosepiece. | A
1.7 | The microscope shall be equipped with an Epi-fluorescence module so the sample can be examined via the oculars and camera. | A
1.8 | The offer shall include a dedicated high-performance anti-vibration table for the confocal microscope, as well as an additional table to support the ancillary equipment required for its operation. | A
1.9 | The system shall include a fully motorized XY stage and shall include a Piezo-Z stage for high-speed volumetric scanning. | A
1.10 | Please specify the travel range, speed, and precision of the Piezo-Z-stage. | I
1.11 | The microscope shall be equipped with several sample holders specifically for 35 mm imaging dishes, multi-wellplates, and slides. | A
1.12 | The system shall have hardware-based Z-drift compensation and software based autofocus. | A
1.13 | The supplier shall offer, as an option, incubation and environmental control, including temperature and humidity regulation, as well as a CO₂ mixer. The option shall be priced separately in Appendix 3.0 – Price Form K935.10, under tab 3 -Utstyr Opsjon. | A
1.14 | The Supplier shall describe additional gas control options, including but not limited to oxygen (O₂) control for hypoxia simulation, if avaliable. Potential options shall be priced separately and listed as individual line items in Appendix 3.0 – Price Form K935.10, under tab 4 - Tilvalg. | A
General Optics and Visual Inspection Modalities
1.15 | The instrument shall include six (6) chromatically corrected APO objectives: 10x Air, 20x Air, 20x Water Dipping, 40x Water Dipping, 60x Oil and 60x Silicon/Glyserol. | A
1.16 | Ref point 1.15. High Numerical Aperture (NA) and long Working Distance (WD) are considered advantageous. The evaluation of this criterion will be based on the user’s assessment of the objectives overall functionality, and performance of the objectives. | B3
1.17 | The vendor shall provide a list of objectives not included in item 1.15, that the vendor think are good objectives for the intended functions/use. These objectives may have different magnification, immersion type, working distance, NA values, etc, than the objectives listed in 1.15. The objectives are to be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 4. Tilvalg. The price is not subject to evaluation. | A
1.18 | The microscope shall be equipped with differential interference contrast (DIC), phase contrast or other equilvalent bright field contrast method. | A
1.19 | The system shall be capable of performing brightfield and DIC (or equivalent contrast) imaging using a reflected-light pathway (from above the sample). | A
1.20 | The microscope shall be equipped with a monochromatic camera optimized for epifluorescence imaging, with a minimum Quantum Efficiency (QE) of 80%, and a color camera for epifluorescence and bright field imaging. | A
1.21 | The microscope should have a large optical field of view (FOV) at the camera port. Please state the optical field of view at the camera port. The microscope with the highest FOV at the camera port will get the highest score, and microscopes with smaller FOV get a lower score, calculated linearly from the maximum. | B2
1.22 | The system should have a large effective field of view. Please specify the maximum achievable Field of View (FOV) using the offered 10x dry objective at an excitation wavelength of 488nm. Evaluation will be based on the diameter (mm) of the usable imaging area. The system with the largest documented usable imaging area will be awarded the highest score. | B3
Excitation and Scanning
1.23 | The system shall include 4 lasers centered at the wavelengths 405nm, 488nm, 561nm, and 647nm, or a single laser that covers the same wavelenghts. | A
1.24 | One additional laser line in the infrared scanning area (e.g. 685nm, 730nm, 750nm or 785nm) shall be offered as an option. The additional laser shall be priced in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Utstyr Opsjon. All additional parts required for full functionality shall be clearly stated and included in the option price. State the wavelenght of the offered infrared laser line. | A
1.25 | The instrument shall have a motorized laser point skanner supporting zoom, bidirectional and ROI scanning. | A
1.26 | The system shall reach acquisition speeds of at least 10 fps at a resolution of 1024 x 1024 pixels. | A
1.27 | Ref. point 1.26, the system will be evaluated based on acquisition speed (frames per second, fps) at a resolution of 1024 × 1024 pixels. An acquisition speed above 10 fps is preferred and will receive a higher score. The system with the fastest acquisition speed will get the higest score. | B3
1.28 | The system shall be able to combine brightfield and confocal simultaneously, and brightfield with super-resolution simultaneously. Describe the offered solutions and imaging speed in the two simultaneous modes. | A
Detection and Resolution
1.29 | The system shall have a minimum of 4 high-QE confocal detectors and a brightfield/DIC detector. | A
1.30 | Ref. point 1.29. Please specify the QE level and the spectral sensitivity. Please attach the Detector Quantum Efficiency and the spectral sensitivity curves for the offered solution. Higher Quatum Efficiency and broader spectral sensitivity will be scored higher. | B3
1.31 | The detector should have a low level of dark noise. Please specify detector dark noise level. Solutions with detectors with the lowest dark noise levels wil be scored the highest. | B1
1.32 | The transmitted light detector (T-PMT/BF detector) must be capable of being positioned (or the system must utilize a reflected-light detector) to allow for the visual inspection and digital capture of opaque samples. | A
1.33 | Please describe the efficiency of the light path (e.g., filter-based or prism-based dispersion). Please emphasize the system's ability to minimize spectral crosstalk through hardware design or software-based spectral unmixing during multi-channel acquisition. | I
1.34 | Detectors shall allow for spectral detection. | A
1.35 | The system shall be able to perform sub-diffraction (super resolution) imaging. The resolution must be better than 250nm. | A
1.36 | The system should provide high spatial resolution for super-resolution imaging. Reference point 1.35. Please describe the specific mechanism/technology for the offered super-resolution. The evaluation will be based on the users’ assessment of the proposed solution and the highest score will be given to the system with the lowest value in nm (x,y resolution). | B2
Advanced Modalities
1.37 | Multiple scanning modes, frame, line, and ROI scan is required. | A
1.38 | The software shall provide full support for FRET, FRAP, and automated tile-stitching. The software must seamlessly stitch the scanned tiles. | A
1.39 | The system shall be equipped with automated navigation tools (such as sample finder or sample overview), including the ability to utilize pre-defined sample templates for standardized multiwell plates. | A
1.40 | The offered solution shall come with a comprehensive software package for advanced data aquisition and analysis. It shall include AI/deep learning functionality for image acquisition, image processing, and image analysis. The Supplier shall give a detailed description of the included functionalities, including, but not limited to, preprocessing, AI-supported analysis (live and post-acquisition), deconvolution/denoising, segmentation, and data analysis. | A | The Supplier shall include user manual(s) for the offered software package. If the Supplier refer to the user manual for describing the functionalities, relevant page numbers must be given for the different functionalities.
1.41 | Reference point 1.40 This criteria looks at the offered software package as described by the Supplier in Point 1.40 and the functionalities thats included. The evaluation will be based on the users’ overall assessment of the functionality of the proposed solution/package. | B1
1.42 | Please describe the systems AI-driven or "smart" acquisition function and if the software can automatically identify regions of interest (ROIs) for high-resolution confocal scanning based on initial widefield detection. Not subject to evaluation. | I
1.43 | Please specify if there are any Optical or software correction for deeper sample scanning included in the solution. Not subject to evaluation. | I
1.44 | Please specify whether the instrument is open for future system upgrades (e.g. adding additional laser lines, new technology). Please also describe the upgrade capability of the system and list potential system upgrade and expansion in the options. Not subject to evaluation. | I
Two-photon options
1.45 | The system shall include an option for a fiber-coupled, femtosecond pulsed laser with dispersion compensation optimized for two-photon excitation of red fluorophores (approximately 1040–1064 nm). The option shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Utstyr Opsjon. The stated price shall include all components, functionality, and deliverables specified in items 1.45 through 1.49. | A
1.46 | Please describe the offered two-photon laser system as specified in point 1.45, including its technical configuration, functionality, output power, and performance characteristics. The system will be evaluated based on the users’ assessment of the overall functionality, stability, power and performance. | B3
1.47 | The offered solution shall include one Non-Descanned Detector 2 (NDD 2) -photon dectector. | A
1.48 | NDD-2 detector with a high QE value is preffered. Please describe the offered NDD 2-photon detector and its QE value, ref. point 1.47. This criteria will be evaluated based on the detectors Quantum Efficiency (QE), where higher QE values will receive higher scores compared to lower QE values. | B2
1.49 | The offered option shall include one 20x/25x water immersion objective that is optimized for deep two-photon imaging. | A
1.50 | The Supplier shall provide a list of optional objectives not included in item 1.49, which the vendor considers suitable for two-photon imaging. These objectives may have different magnification, working distance, NA values etc, than the objective described in 1.48. The listed objectives shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 4. Tilvalg. Prices will not be subject to evaluation. | A
1.51 | The system should utilize an optical design that minimizes the optical path length between the sample and the NDD detector. Please specify the optical design and the physical placement of the NDDs relative to the objective lens in in millimeters or centimeters from the back aperture of the objective to the NDD sensor. The detector configuration will be evaluated based on the physical optical path length between the sample and the detectors, where shorter optical paths will receive higher scores compared to longer optical paths. | B3
1.52 | The two photon option should include the possibility for Fluorescence Lifetime Imaging Microscopy (FLIM). If FLIM is supported, please describe the underlying technology (time-domain or frequency-domain) and specify the acheivable lifetime measurement range. The proposed solution will be evaluated based on whether FLIM functionality is possible with the two-photon option. If available, the FLIM solution shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 4 - Tilvalg. The stated price must encompass all necessary hardware, detectors, electronics, and software required for a fully functional workflow. | B1
General Requirements
1.53 | The instrument and its associated transport packaging must fit through a door of 90x205cm (WxH). Describe the number, dimensions and weight of the transport packages. | A
1.54 | Please describe how the Supplier handles product updates or new software releases introduced between the time of tender submission and the actual delivery of the instrument. Please state whether upgrades to newer versions are included in the offered price, or if they entail additional costs. This point is not evaluated. | I
IT Hardware
1.55 | The offered instrument shall include a dedicated workstation with a high-performance graphics card, high performance storage drives, a keyboard, and a mouse. The instrument computer shall be delivered with all software modules necessary for instrument control. | A
1.56 | Two high-resolution monitors (27" or larger), or a single ultra-wide monitor providing an equivalent total screen area and display resolution shall be included in the offered solution. The Tenderer shall provide a technical description of the monitor(s) included with the instrument, explicitly specifying the screen size and display resolution. | A
References
1.57 | References. List at least two relevant references for the offered solution. The references shall relate to comparable installations and shall include contact details such as telephone number, e-mail address, and workplace. | A
Demonstration
1.58 | The supplier shall offer a demonstration of the offered solution, to be held during the evaluation phase of the tender. Please see point 4.4.1 in document "1.0 Invitasjon til anbudskonkurranse K935.10" for additional information on demonstrations, including relevant time table. It must be a physical demonstration at a suitable location. | A
1.59 | The demonstration of the offered solution, should be with the instrument set-up as described in this tender. The buyer have the right to use their own samples for the demonstration, should they wish it. The evaluation will be based on the users’ assessment of the proposed solution as shown in the demonstration, including user interface. Suppliers should showcase the following capabilities: - Live Sample Imaging: Water-dipping imaging of a live sample using at least two fluorophores, combined with brightfield and multi-point (XYZT) acquisition. - 3D Deep-Tissue Scanning: A 3D Z-stack of a thick sample to demonstrate the effective working distance and performance of the water-dipping objectives. - Large-Area Tile Scanning: A multi-channel tile scan of a large sample (minimum 20 fields of view) to demonstrate stitching quality and illumination uniformity. - Super-Resolution Mode: A demonstration of the system's super-resolution capabilities in terms of spatial detail and acquisition speed. - Image Processing Software: A demonstration of the software suite, specifically focusing on data handling, visualization, and analysis. -Optional: 2-Photon and FLIM: Performance demonstration of the 2-Photon excitation depth and the Fluorescence Lifetime Imaging (FLIM) module is preffered but not required. If a physical demonstration of any of these modes on the actual systems is not possible, a powerpoint (or similar) presentation containing technical details of these modes must be included instead. | B3
Training requirements
1.60 | Necessary super user training must be offered for approx. 5 users. The training must be carried out by agreement. The traning is done on-site at the customer site. The training must ensure proper, safe and correct use of the equipment. The Vendor must anticipate that training will take place after takeover. Describe scope and content of the training. Training must be included in the tender at no additional costs. | A
1.61 | Please describe the capability and availability for providing on-demand training sessions for new or advanced users after the initial system installation period. | I
ICT requirement
1.62 | All mandatory program and security updates for both software and equipment shall be delivered and installed free of charge until the end of service. By update is meant improvement and "bugfix" of software. But not new features. | A
2 | SERVICE. OPERATION AND MAINTENANCE CRITERIA
Service requirement
2.1 | The Vendor shall list prices for the following in Appendix 3.0 –Prisskjema K935.10 under tab 3. SERVICETJENESTER : - Attendance price, the price shall include all travel expenses and time spent per travel to the service point - Hourly price for the performance of service, covers time spent with the client or supplier These rates/prices are included to the evaluation of total cost in the tender. | A
Maintenance agreement
2.2 | An option for a maintenance agreement shall be offered Level 1: Preventive The maintenance agreement shall include: - Preventive maintenance according to the manufacturer's specifications, including all wear parts/service kits/components/batteries that the manufacturer recommends replacing in connection with maintenance, including parts, all travel and accommodation expenses and any application follow-up or similar. - Documentation of work carried out (service report) shall be delivered to the client within 5 working days - Phone support The supplier shall schedule in advance the time for the maintanance with the Buyer. The agreement is to be priced in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Vedlikeholdsavtaler opsjon | A
2.3 | An option for a maintenance agreement shall be offered. Level 2: Comprehensive - Comprehensive agreement with annual costs for preventive maintenance including parts, emergency calls, including parts, all travel and accommodation expenses and any application follow-up or similar. The agreement shall cover all breakdown repairs and parts, including lasers. - Documentation of work carried out (service report) shall be delivered to the client within 5 working days - Phone support The supplier shall schedule in advance the time for the maintanance with the Buyer. The agreement is to be priced in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Vedlikeholdsavtaler opsjon. | A
Service and maintenance
2.4 | Please state the maximum delivery time from order to delivery at the customer's place, including delivery from "external warehouse". (1 calendar day = 24 hours). | I
2.5 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual shall be in English or a Scandinavian language. | A
2.6 | Phone support shall be available during normal working hours. Communication with customer support shall be carried out in a Scandinavian language or English. Ordinary working hours are defined as weekdays at 08:00-16:00. Describe the support available and any conditions or limitations to this service. | A
2.7 | Please give a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for carrying out preventive maintenance, as well as information on frequency and expected time consumption should be attached. | I
2.8 | State the offered warranty period for the system. | A
2.9 | During the warranty period, the vendor must provide preventive maintenance at service level 1 (see point 2.2), at no additional cost to the buyer. Periodic maintenance must be carried out in accordance with the manufacturer's recommendations. The final periodic maintenance must be performed within the two last month before the warranty period expires. | A
2.10 | Buyer request an option for an Extended Warranty Agreement. Note point 2.8. The agreement shall be priced in 3.0 Prisskjema K935.10 under tab 3 - Garanti opsjon. | A
2.11 | At the end of the warranty period, delivered equipment shall have the latest updated version of the software. There shall be no additional costs for this. | A
2.12 | If there is a need for equipment-specific special equipment, calibration tools, software, passwords, service keys, or similar, to carry out technical troubleshooting, maintenance and/or emergency repairs, this must be explained and priced in Appendix 3.0 – Prisskjema K935.10 under tab 4. Tilvalg | A
2.13 | The Supplier shall describe which parts are expected to be replaced during the equipment's expected lifetime, and the frequency of replacement. E.g. light sources, moving parts, other? | A
2.14 | The Vendor must guarantee access to spare parts, at least 7 years after delivery of the equipment. | A
2.15 | Access to spare parts over 7 years is emphasised. State how many years in total the provider guarantees access to spare parts. | B3

## 3. Technical specifications
TECHNICAL SPECIFICATIONS
Statsbygg art. nr | Article name | Tenderers offer/model | Acoustic noise: (dB (A)) | Heat emission: Standby / operation (kW) | Room requirements: Temperature (oC incl peak-to-peak limits) Relative humidity min/max (% rH) | Power supply: (power plug, fixed plug, length of power cord) Effect (kW) Voltage (V) Ampere (A) Require dedicated circuit Require dedicated earthing | Specific electricity requirements: Emergency power supply UPS | Ethernet: Socket position, cable type, cable routing, plug type, other, capacity Any max cable length | Lighting source Cable routing | Other:
025.02.001:3 | Konfokalmikroskop, Point Scanning Confocal
Statsbygg art. nr | Article name | Tenderers offer/model | Couplings to: Gases (incl compressed air) Pressure (bar) Purity (%) | Other:
025.02.001:3 | Konfokalmikroskop, Point Scanning Confocal