## Explanation
Explanation for the requirements
White field must be filled in by supplier.
Supplier must not add or delete rows or columns in the requirements.
Column | Explanation of requirements specification
Referance | Buyer has given each row a unique identification number. This should make it easier to refer to the correct point in the requirements specification.
Description of requirements | Buyer has described different requirements related to the product in form of minimum requirements, desired requirements and information requirements.
Type of requirement | Minimum requirements (A) Requirements are important and must/shall be met. If not, Suppliers offer will be rejected. Desired requirements (B) Requirements should be met and should give added value for the buyer. This is not an absolute requirement and supplier will not be rejected if they do not fullfill this requirement. Answer will be subject to asessment.
Suppliers answer: Does supplier meet requirements? | Supplier must tick X to show that their product meets requirements.
Check X for YES = requirements met.
Check X for NO = Requirements not met.
Suppliers answer: Detailed description | Supplier must fill in a detailed description where it is necessary. Supplier should give complementary information that can enlight the requirement.
Type of documentation: | Document: Suppliers documentation in form of attachment, link or anything that can substantiate suppliers answer to question. Answer to question and how supplier provide buyer with documentation can be part of an asessment.

## Requirements
Specific requirements
| Buyers requirements |  |  | Supplers answer
Ref. | Description of requirement | Type of documentatation | Type of requirement | Requirement fullfilled? |  | Description
|  |  |  | yes | no
1.0 | Core Instrument Performance and Specifications
1.1 | The SEM shall achieve a minimum imaging resolution of 0.6–0.7 nm or better under standard operating conditions. More details: The supplier shall confirm that the instrument meets this resolution specification. Supporting documentation such as a datasheet or factory performance certificate shall be uploaded. | Document | A
1.2 | Electron source shall meet the performance requirements: i) Resolution at ≤ 1 kV ≤ 0.8 nm; ii) Resolution at ≤ 15 kV ≤ 0.7 nm; iii) Resolution at ≤ 30 kV ≤ 0.6 nm. More details: All three resolution thresholds must be met as minimum requirements. | Document | A
1.3 | Resolution at ≤ 1 kV — the actual resolution achieved is weighted in the evaluation. More details: Provide the best achievable resolution at ≤ 1 kV accelerating voltage. Superior performance beyond the minimum threshold will be evaluated positively. | Number | B
1.4 | Resolution at ≤ 15 kV — the actual resolution achieved is weighted in the evaluation. More details: Provide the best achievable resolution at ≤ 15 kV accelerating voltage. Superior performance beyond the minimum threshold will be evaluated positively. | Number | B
1.5 | Resolution at ≤ 30 kV — the actual resolution achieved is weighted in the evaluation. More details: Provide the best achievable resolution at ≤ 30 kV accelerating voltage. Superior performance beyond the minimum threshold will be evaluated positively. | Number | B
1.6 | The instrument must be equipped with a cold or Schottky field emission gun (FEG) electron source More details: The supplier shall confirm that the instrument features a cold or Schottky field emission gun as the electron source. | Describe | A
1.7 | The SEM should be equipped with a Schottky FEG or Cold FEG for robust operation and sufficient sample current. The specifications in terms of current, lifetime, and cost of replacement are weighted in the evaluation More details: Supplier should describe the electron source type, typical emission current, expected tip lifetime, and estimated replacement cost. | Describe | B
1.8 | The instrument shall support a wide accelerating voltage range suitable for both beam-sensitive and conductive materials. More details: Specify the full accelerating voltage range available on the instrument (in kV, from minimum to maximum). A wider range enabling low-voltage operation is advantageous for imaging beam-sensitive and non-conductive samples. | Describe | A
1.9 | Features to mitigate beam damage of beam-sensitive samples should be provided and will be weighted in the evaluation. More details:Describe available features, scan modes, or accessories that reduce beam damage, such as low-kV operation, beam blanker, or other protective mechanisms. | Describe | B
1.10 | The microscope system shall be equipped with at least 1 detector for Secondary Electrons (SE), 1 detector for Backscatter Electrons (BSE), 1 STEM detector and 1 Energy Dispersive X-ray Spectroscopy (EDS) chemical analysis system. |  | A
1.11 | The instrument shall include a BF-STEM (bright-field scanning transmission electron microscopy) detector as part of the standard delivery. More details: The supplier shall confirm that a BF-STEM detector is included in the delivery. This is required for transmission-mode imaging of nanoparticles and thin specimens. | Describe | A
1.12 | Modes of STEM detector operation - i.e., generating contrast from phase, diffraction, and atomic number as well as STEM sample holder features should be provided. | Describe | B
1.13 | The instrument shall include an EDS (energy-dispersive X-ray spectroscopy) detector for chemical characterisation. System must include seamless operation with SEM (hardware and software), and document time needed to switch between high res -SEM imaging and EDS conditions. More details: The supplier shall confirm that an EDS detector is included in the standard delivery. Supporting documentation (e.g. detector specifications) may be uploaded. |  | A
1.14 | One or more EDS solutions should be offered, including both a general-purpose solution and solutions optimized for the analysis of nanostructures. The supplier shall confirm the proposed solutions and may provide relevant specifications. |  | B
1.15 | Number, performance, and integration of SE and BSE detectors, including electron enegy filtering shoul be provided. More details: The number of detectors for different modes of operation and detection and their performance and integration with software and other hardware is important. | Describe | B
2.0 | Sample Handling and Accessories
2.1 | Plasma cleaner installed on main chamber or airlock must be provided. More details: A plasma cleaner must be included as part of the delivery, mounted on the chamber or airlock. |  | A
2.2 | Integration and operation of plasma cleaner in SEM software should be provided More details: Provide documentation describing how the plasma cleaner is integrated and controlled via the SEM software. |  | B
2.3 | Cold finger on chamber for specimen anti-contamination shall be provided More details: A cold finger or anti-contamination device must be included on the chamber. |  | A
2.4 | System with airlock should be provided More details: Describe whether an airlock is included in the delivery and its specifications. | Describe | B
2.5 | Inert sample transfer of samples prepared for SE/BSE/STEM and EDS analysis using an inert transfer shuttle between glovebox and SEM. More details: An inert transfer shuttle enabling fully air-free transfer from a glovebox to the SEM shall be delivered | Describe | A
2.6 | Integration and operation of the inert sample transfer shuttle unit with respect to handling in glovebox (dimensions of transfer unit to fit in small transfer chamber in the glovebox) and duration of keeping the sample inert should be delivered and will be weighted in the evaluation. MOre details: Provide documentation on the physical dimensions of the transfer unit, compatibility with standard glovebox transfer chambers, and the duration for which inert atmosphere is maintained. | Describe | B
2.7 | Sample stage capabilities and control — stage capabilities, controls and automation, and systems to prevent crashes are important and shall be met. |  | A
2.8 | Ports for future additions on airlock and main chamber — possibilities for extra add-ons in the future, seen in view of the components added during delivery, should be provided More details: Provide documentation describing available ports and upgrade possibilities on both the airlock and the main chamber. | Describe | B
3.0 | Software, Imaging, and Data Acquisition
3.1 | The instrument should include a camera-based navigation system (overview camera or stage navigation camera) enabling efficient sample navigation. More details: The supplier should confirm that a camera navigation system is included in the delivery, enabling operators to locate and navigate to regions of interest efficiently. |  | B
3.2 | Describe the software capabilities for EDS data acquisition, mapping, and quantification. More detail: Provide a description of the software tools available for EDS analysis, including elemental mapping, quantification routines, and data export formats. More comprehensive and user-friendly EDS software will score higher in the evaluation. | Describe | B
3.3 | Describe the user interface and automation features available for multi-user and multi-level operation (BSc, MSc, PhD, and senior researcher levels). More details: Provide a description of how the instrument's software and interface supports users of varying expertise levels, including any automation, guided workflows, or user profile management features. Systems offering intuitive interfaces and automation for less experienced users will score higher. | Describe | B
3.4 | Computer system and software — the computer system with all required software must be included. More details: A complete, operational computer system controlling the microscope and all attached hardware must be included in the delivery. Any items to be procured separately must be clearly identified. |  | A
3.5 | Computer system and software capabilities, performance, integration, and user interface — the computers and software that control and provide output of the microscope and all supplied attached hardware should be functional and capable as a whole. Particularly, intuitive and pedagogical software promoting fast learning of the principles will be weighted in the evaluation. More details: Describe the user interface, software architecture, and any features that facilitate rapid user onboarding and multi-level operation. | Describe | B
3.6 | Automated sample analysis — the microscope will serve a wide range of user patterns from routine research investigations to methodological frontier studies, so automated sample investigations (e.g., several samples or areas) are considered useful but at moderate priority. More details: Describe any automated analysis features available, such as automated area scanning, particle analysis, or batch processing. | Describe | B
3.7 | Stand-alone software for post-processing of analyses and images will be awarded. Describe the offered software and number of licenses included in the bid. Pricing for additional licenses shall be provided in the price schedule. More details: Describe any stand-alone post-processing software included, the number of licenses, and indicate the price for additional licenses in the pricing schedule. | Describe | B
3.8 | All mandatory hardware, middleware, software and security updates must be provided and installed without cost until end of service (updates are to be interpreted as necessary improvements and bug fixes). More details: The supplier must provide all necessary updates for the lifetime of the instrument at no additional charge. |  | A
4.0 | Installation, Training, and Service
4.1 | The supplier shall provide full on-site installation and commissioning at the University of Oslo's premises. More details: The supplier shall confirm that full installation and commissioning, including performance verification, is included in the delivery and will be carried out at the University of Oslo's Department of Chemistry. |  | A
4.2 | The supplier shall provide an on-site user training programme covering all relevant user levels (BSc, MSc, PhD, and senior researchers). More details: The supplier shall confirm that an on-site training programme tailored to users at different levels of expertise is included in the delivery. This includes also super users. Please describe the content, duration, and structure of the training programme. | Describe | A
4.3 | The supplier should include a comprehensive service agreement covering preventive maintenance and remote diagnostics. More details: The supplier should confirm that a service agreement including scheduled preventive maintenance and remote diagnostics capability is included as an option or offered as part of the delivery. Documentation of the service agreement terms may be uploaded. | Describe | B
4.4 | The supplier shall provide full documentation of infrastructure requirements prior to delivery, including power supply, water cooling, air-conditioning, floor stability, need for electromagnetic radiation cancelling system and possible vibration isolation specifications. More details: The supplier shall confirm that complete infrastructure requirement documentation will be provided to allow the University of Oslo to make all necessary preparations. The instrument will be installed in EM area in new life science building ( Support room for pumps and chillers) |  | A
4.5 | Supplier should inform about the estimated delivery time from reception of order to the system is delivered to site of installation. If different parts of the system offered have different delivery times, this is to be stated. More details: Provide the expected date by which the instrument will be fully installed and commissioned at the University of Oslo's premises. State the expected duration of the on-site installation process. |  | B
4.6 | Electrical power requirements — the building has 3-phase 400 V AC as standard electrical power supply. If the instrument can run on this it is considered an advantage. Alternatively, 1-phase230 V AC is an option. More details: State the electrical power supply required for the instrument. Compatibility with 3-phase 400 V AC will be weighted positively. | Describe | A
4.7 | The system must be delivered with 2 years warranty. More details: Confirm that the offered system includes a minimum 2-year warranty. Provide written confirmation. |  | A
4.8 | Spare parts must be available from the supplier for at least 10 years after the procurement date. More details: Describe the supplier's commitment and policy for spare parts availability over the 10-year period. | Describe | A
4.9 | The supplier should specify all spare parts, consumables, and sample holders included in the standard SEM delivery. The standard delivery shall include all components necessary for routine operation and installation of the system. The supplier may offer additional sample holders and spare parts for specialized applications and shall provide specifications for the proposed items. | Describe | B
5.0 | Environmental and Sustainability Considerations
5.1 | Supplier should have an approach to end-of-life management and take-back of the instrument. More details: Provide information on whether the supplier offers take-back or recycling programmes for the instrument at end of life, and how hazardous materials and components are handled. | Describe | B
5.2 | Instrument life time must be at least 10 years. |  | A
5.3 | Supplier must specify any parts that must be replaced withing the first 5 years and have a cost exceeding 2.5% of the instruments purchase price |  | A
6.0 | Product Requirements – Compliance, Documentation, and Pre-installation
6.1 | The equipment shall be CE marked in accordance with all relevant directives and comply with all relevant Norwegian laws, regulations, standards, requirements and norms for the equipment offered. More details: Declaration of conformity shall be attached with the bid. |  | A
6.2 | Complete user manual(s) and safety data sheet(s) in a Scandinavian or English language shall follow the delivery, preferably in electronic formats. The supplier shall, until end of service as given by the manufacturer(s), provide unsolicited electronic updates of the user manual(s) and safety data sheet(s), regardless of file format. More details: Manuals and safety data sheets must be kept up to date throughout the service life of the instrument. |  | A
6.3 | Supplier shall deliver quote, brochure, datasheet and all neccessary documents in a structured and orderly manner so that the contracting authorithy can easily navigate in the tender. Note: This spreadsheet must not be converted to PDF. |  | A
6.4 | The supplier shall, if relevant, conduct site survey(s) in due time before installation in order to ensure that all installation site requirements are in accordance with the manufacturer's requirements. All costs related to site survey(s) shall be included in the bid. More details: Site surveys should cover vibrations, electromagnetic radiation, power supply, water cooling, compressed air, and requirements for temperature/humidity control in the room. |  | A
7.0 | Optical and Imaging Modes
7.1 | The microscope should have an optical camera for surveillance of the chamber and sample stage. Describe the solution offered. More details: An optical/chamber camera should be included for stage and sample monitoring. | Describe | B
7.2 | Solutions including an optical imaging mode should be provided | Describe | B
7.3 | Solutions including a chamber TV camera should be provided More details: Describe whether a dedicated TV camera for the chamber is included or available. | Describe | B
Supplier must not add or delete rows or columns in the requirements. Do not convert this document to PDF or other formats. Drawings and documentation must be enclosed

## Online demo requirements (hidden)
Specific requirements
Buyers requirements
Ref. | Description of requirement | Type of documentatation | Type of requirement
1.0 | Instruments and components
1.1 |  | Online demo | B
1.2 |  | Online demo | B
1.3 |  | Online demo | B
1.4 |  | Online demo | B
Software
|  | Online demo | B
|  | Online demo | B
|  | Online demo | B
|  | Online demo | B
|  | Online demo | B
This sheet shows our focus pint on online demo and shall not be answered by supplier

## Price scheme
Supplier | EX.VAT.
Price equipment
Number of units | Descripion | Part number | Price pr.unit | Sum
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
Sum |  |  |  | 0
Price service and maintenance (incl.parts,travelling,hotels, etc) after warranty period
Number of units | Desciption |  | Price pr.unit | Sum
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
Sum |  |  |  | 0
Life cycle cost spare parts after warranty period (2 years)
Number of units | Description |  | Price pr.unit | Sum
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
Sum |  |  |  | 0
Options
Number of units | Description |  | Price pr.unit | Sum
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
|  |  | 0 | 0
Sum |  |  |  | 0