Appendix 1: Customer Requirements Specification

# 1. About the Customer

The Customer is the University of Bergen (UiB). UiB is a medium-sized European research university with approximately 20,000 students and more than 4,000 employees. This procurement is carried out on behalf of the Department of Biological Sciences (BIO). The instrument will support research in marine microbiology, environmental microbiology, ecology, biotechnology and molecular biology, and will form part of a developing flow cytometry core facility serving several research groups.

# 2. Purpose of the Procurement

The purpose of this procurement is to acquire a modern, versatile and user-friendly flow cytometer capable of analysing a wide range of biological particles and cells relevant in marine and environmental research, biotechnology, and molecular biology. The instrument shall support analysis of particulate biological material, including viruses, bacteria, extracellular vesicles, and single eukaryotic cells in the size range <100 nm – 50 µm.

The instrument will be installed in a dedicated laboratory and operated in a multi-user core facility environment. Robust, intuitive workflows are therefore required to ensure efficient operation by users with varying levels of experience, while allowing advanced users to develop specialised applications.

The acquisition is financed by the Faculty of Science and Technology, UiB. Maximum budget: NOK 1.76 million (VAT excl.). This budget shall include a minimum two-year warranty agreement for the complete system, including any subcontracted components.

Within this budget frame, suppliers are invited to propose the most versatile high-end flow cytometer system possible, meeting high demands for performance, flexibility and usability, and fulfilling the requirements in this specification.

Options will not be included in the maximum budget unless explicitly stated but may be included in the evaluation.

# 3. Scope of Delivery

The Supplier shall deliver a complete and fully operational flow cytometer system, including all components and services necessary for its intended use.

The scope of delivery shall include, but not be limited to:

- Flow cytometer instrument

- Control computer and all required hardware

- Data acquisition software

- Data analysis software (including licensing as specified)

- Installation, commissioning and functional verification

- On-site user training (minimum as specified)

- Documentation (user manuals, service/maintenance documentation, and compliance matrix)

- Warranty, service and support (minimum 24 months)

The delivered system shall be ready for use upon completion of installation and acceptance testing. Acceptance testing shall verify that the system meets the agreed functional and performance requirements.

# 4. Requirements

This section specifies the Customer’s Requirements regarding functionality, performance and interfaces with which the delivery must comply.

The Supplier shall provide a detailed description of how the offered solution meets each requirement and shall document the performance and relevant parameters of the instrument by providing supporting data and/or references to published information. The objective is to ensure a common understanding between the Customer and the Supplier of what is to be delivered.

It is not expected that the Supplier will meet all requirements in full. However, it is a requirement that any deviations or exceptions are clearly described in the Supplier’s solution specification. Where a requirement can be met only with reduced or non-optimized performance, the consequences for system performance shall be clearly explained.

The requirements are assigned different priorities:

- Priority A – Main requirements that are important for the intended use. Responses will be significant in the evaluation and are weighted higher than Priority B requirements.

- Priority B – Important requirements that contribute to the overall quality of the solution. Responses will be considered in the evaluation and are weighted lower than Priority A requirements.

In the evaluation, Priority A requirements will be weighted higher than Priority B requirements.

The Supplier shall provide references to relevant documentation and include any necessary comments in the response. It is the responsibility of the Supplier to document compliance with the stated requirements.

Where requirements request descriptions, evaluation will emphasize clarity, level of detail, documented performance, and relevance to intended use.

# 4.1 Hardware and Instrument Configuration

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| HW-01 | A | The system offered must be a fully integrated flow cytometer designed for research applications. |  |
| HW-02 | A | The instrument must support detection and analysis of biological particles across a broad size range, including particles smaller than 100 nm and particles up to at least 50 µm in size. |  |
| HW-03 | A | The system must support analysis of small biological particles such as bacteria and viruses. |  |
| HW-04 | A | The Supplier must specify the expected operational lifetime and support period (target ≥10 years). |  |
| HW-05 | B | The system should be designed for stable operation in multi-user laboratories. |  |
| HW-06 | B | The instrument should support long measurement runs with minimal drift. |  |
| HW-07 | A | The Supplier must ensure availability of spare parts and service support for a minimum of 10 years from the date of final acceptance. |  |

# 4.2 Optical Configuration

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| OPT-01 | A | The instrument must include blue, violet, yellow-green and red lasers (approximately 405 nm, 488 nm, 561 nm and 638 nm), or equivalent configurations providing comparable performance. |  |
| OPT-02 | A | The system must support simultaneous detection of multiple fluorescence signals. |  |
| OPT-03 | B | The system should allow future installation of additional lasers or detectors. |  |
| OPT-04 | B | The Supplier should specify optical configuration and detector types. |  |

# 4.3 Scatter Detection (FSC / SSC)

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| SC-01 | A | The instrument must include forward scatter (FSC) and side scatter (SSC) detection. |  |
| SC-02 | A | The instrument must provide accurate and precise side scatter (SSC) detection for small biological particles (<100 nm). This can be achieved by the presence of an additional laser at approximately 405 nm or equivalent technical solutions providing comparable performance. |  |
| SC-03 | A | SSC detection must support analysis of small biological particles such as bacteria and viruses. |  |
| SC-04 | B | The Supplier should describe how scatter detection is optimized for small particles. |  |
| SC-05 | B | Scatter detector response should be linear over the stated dynamic range. |  |

# 4.4 Fluorescence Detection Capability

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| FL-01 | A | The system must support multi‑parameter fluorescence detection and analysis. Systems supporting ≥6 fluorescence parameters are preferred. |  |
| FL-02 | B | The Supplier should describe how the system addresses spectral overlap (compensation, spectral unmixing or equivalent). |  |
| FL-03 | B | Fluorescence detector response should be linear over the stated dynamic range. |  |

# 4.5 Analytical Performance

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| PERF-01 | A | The system must provide high optical sensitivity suitable for small biological particles. |  |
| PERF-02 | A | The Supplier must demonstrate or document detection limits using experimental data. |  |
| PERF-03 | B | The system should provide high throughput suitable for environmental sample analysis. |  |
| PERF-04 | B | The Supplier should describe signal-to-noise performance. |  |

# 4.6 Sample Handling

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| SMP-01 | A | The instrument must support standard sample tubes. |  |
| SMP-02 | B | The instrument should support connection of automated sample loaders for both tubes and 96-well plates. |  |
| SMP-03 | B | The Supplier should specify the expected sample throughput under typical operating conditions and describe factors affecting throughput. |  |

# 4.7 Software

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| SW-01 | A | The analysis software must be capable of running on a standard high‑performance Windows PC (not connected to the instrument) or equivalent architecture that enables offline analysis. |  |
| SW-02 | A | All software included in the offer must be specified (acquisition, analysis, drivers, libraries). |  |
| SW-03 | A | The software must support and be able to export data in open standard file formats, e.g. FCS 3.0 and 3.1. |  |
| SW-04 | A | The software must support gating, statistics and data visualization. |  |
| SW-05 | B | The software should be user-friendly, including intuitive workflows, documentation, and suitability for multi-user environments. |  |
| SW-06 | B | The Supplier should describe options, policy and potential costs for future software upgrades. |  |
| SW-07 | B | An unlimited number of analysis licenses is preferred. |  |

# 4.8. Delivery

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| DEL-01 | A | Delivery terms: Incoterms 2020 DDP VAT unpaid, delivered to the Department of Biological Sciences, University of Bergen, Bergen, Norway. |  |
| DEL-02 | A | All delivery costs must be included in the tender. |  |
| DEL-03 | B | The Supplier should specify delivery time. |  |
| DEL-04 | B | The Supplier should minimize packaging materials and use recyclable packaging where practical. |  |

# 4.9. Installation

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| INST-01 | A | The Supplier must perform installation and commissioning. |  |
| INST-02 | A | All requirements for utilities and laboratory conditions must be specified. |  |

# 4.10. Training

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| TRN-01 | A | On-site user training must be included. |  |
| TRN-02 | A | Training costs must be included in the tender. |  |
| TRN-03 | B | The Supplier should include advanced training for at least 1–2 core facility super‑users. |  |

# 4.11. Warranty, Support and Maintenance

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| SER-01 | A | The Supplier must clearly specify warranty conditions and included services. |  |
| SER-02 | A | The Supplier must provide a minimum warranty period of 24 months from final acceptance of the system, including service and technical support. |  |
| SER-03 | B | The Supplier should specify service response times and availability of local/remote support. |  |
| SER-04 | B | The Supplier should describe preventive maintenance schedule and typical annual running costs (consumables, service). |  |
| SER-05 | A | The Supplier must provide a documented preventive maintenance programme for the offered system. |  |
| SER-06 | A | The Supplier must describe how the system is designed to support repairability and long operational lifetime, including availability of replacement parts, service personnel and software support. |  |

# 4.12 Options

Optional components may be taken into consideration in the qualitative evaluation where they contribute to increased functionality, flexibility, or future usability of the system. However, the evaluation will primarily emphasize the performance and functionality of the base configuration within the defined budget.

| Req. | Priority | Description | Supplier’s response |
| --- | --- | --- | --- |
| OPTN-01 | A | All optional components must be priced separately and clearly linked to the base configuration. |  |
| OPTN-02 | B | Autosampler should be offered as an option. |  |
| OPTN-03 | B | Spectral detection capability may be offered as an option. |  |
| OPTN-04 | B | Additional software licenses / dongles should be offered as options. |  |
| OPTN-05 | B | Additional annual maintenance service agreements should be offered as options. |  |