## Guidance for use
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Project 1105601 UiO Livsvitenskap brukerutstyr
K946.01.01 Pharmaceutical tablet production equipment
Tenderer:
HOW TO FILL OUT THE DOCUMENT
Please note that there are several tabs to be filled out.
Columns and rows are not to be deleted.
The tenderer shall only fill out fields with a white background. Fields of other colors are not to be used or altered by the Tenderer.
The tenderer shall answer YES or NO to the question on whether or not a requirement is fulfilled. All rows must be answered.
The tenderer shall answer the defined requirements. The answer shall have sufficient information for the contractor to carry out a full assessment of the tender without the need to open attachments. If the tenderer needs to provide additional or more detailed information than this document allows for, it is allowed to refer to attachments. In case of reference to attachments, the tenderer shall provide the corresponding attachment nr and relevant pages.
Minimum requirements (A-criteria) are not related to an award criterion. Lack of fulfillment of minimum requirement can lead to rejecton of the tender, cf Tilbudsinvitasjon (Tender invitation).
B-criteria are related to an award criterion. The tenderers answer to B-criteria are the basis of the contractors evaluation and assessment. B-criteria are defined as B1, B2 or B3 criteria in the corresponding column for each criterion. The definitions indicate the assigned weight of the award criterion in the tender evaluation. Award criteria indicated as B3 are assigned most weight, while the B1 criteria are assigned the least weight.
Information criteria (I-criteria) are solely for informative purposes, and are not part of the evaluation.
| Valuta: | Lisens:
| NOK | Systemlisens
| SEK | Brukerlisens
| DKK | Utstyrslisens
| EUR | Analyselisens
| USD | Annet
GBP

## Subcontract 1 Drum coater
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 029.03.050 Drum coater
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | General requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Intended use: The offered equipment shall be a drum film-coater, used to perform small-scale pharmaceutical tablet coating operations with polymeric coatings for research and supervised student practicals. | A
1.5 | The drum film-coater should process batch sizes up to approximately 1 kg. Please specify the maximum processable batch size. Suppliers with higher capacity (up to or above 1 kg) will be scored higher, provided coating performance is maintained. | B1
1.6 | The drum film coater should be capable of processing batches down to approximately 75 g. Please specify the minimum processable batch size. The ability to process batch sizes close to 75 g will be considered an advantage and will receive a higher evaluation score. | B2
1.7 | If multiple drum sizes are needed to meet the nominal batch-size ranges indicated in Sections 1.6 and 1.7, the supplier should provide the geometric drum volume and the recommended operational working range for each drum. These data will be used to evaluate the suitability and flexibility of the proposed equipment. | B2
1.8 | Uniformity of film thickness should be ensured. Please specify the methods or technologies used to achieve uniform coating. Additionally, please provide the achievable film thickness ranges along with the corresponding distribution or variation for each range. This information will be used to evaluate the coating performance and the suitability of the equipment. | B3
1.9 | The internal drum shall be made of stainless steel. | A
1.10 | The drum speed shall be adjustable. The speed range shall be specified and document in rpm. | A
1.11 | The offered minimum–maximum speed range will be evaluated, and wider ranges will receive a higher score. | B1
1.12 | The spray system shall have adjustable flow rate and atomization control. Please specify how this is achieved. | A
1.13 | Solutions that are assessed to have wider control ranges (flow or atomization) and more precise control methods will receive a higher score, cf. requirement 1.12 | B1
1.14 | The process temperature within the coater should be accurately controlled. Please describe the temperature control system, including the accuracy of the set temperature. The solution will be evaluated based on the described temperature control performance and suitability for the intended application. | B2
1.15 | The inlet air temperature should be controllable over a range suitable for using both aqueous and ethanol-based solutions (ca. 30 to 70 °C). Please describe the control mechanisms and the achievable ranges of the inlet air temperature and flow rate, including the possibility to monitor the temperatures. The ability to monitor both air flow rate and temperature as well as the controllable range of temperature is considered an advantage and will be evaluated based on the described functionality. | B2
1.16 | For the inlet air, please specify the suitable range of humidities of the inlet air in percentage relative humidity and describe whether it is possible to draw the inlet air from the room the equipment is operated in. | I
1.17 | For the inlet air, please specify whether a filter is connected to the inlet air draw such as to ensure the necessary low-particle quality of the inlet air. If no such filter is connected, please describe the suitable particulate ranges for the inlet air to be supplied to the equipment during operation. | I
1.18 | For the outlet air, please specify the maximum necessary air to be exhausted during operation in Liters per minute. This will be used to access whether the equipment needs to be connected to a ventilation channel for exhaust. In addition, please describe how the system can be connected to an exhaust outlet and what type of connector to be used. | I
1.19 | The coating process should be visually observable during operation to be usable for teaching. Good visibility of the coating process is considered an advantage and will be evaluated based on the described solution for visual monitoring. | B2
1.20 | The nozzle spray should be uniform and well-dispersed. Please describe the methods used to achieve this, and specify the available nozzle shapes and sizes. The ability to achieve uniform and well-dispersed spraying is considered an advantage and will be evaluated based on the described spray characteristics. | B1
1.21 | It should be possible to use various coating media, including aqueous and ethanol-based solutions. Please describe the range of solvents that can be used. The ability to use a wider range of coating media is considered advantageous and will receive a higher evaluation. | B2
1.22 | The Eequipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.23 | The supplier shall complete the sheet Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Training requirement
2.1 | The necessary superuser training shall be provided for approximately 2-3 users. The training shall be conducted by agreement and on-site. The training should ensure safe, secure, and correct use of the equipment as well as correct maintenance to ensure longevity of the equipment. The supplier should anticipate that training will take place after handover. The scope and content of the training shall be described. | A
3.0 | Hygiene/Cleaning requirements
3.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
3.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
3.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
4.0 | IT-Requirement
4.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
5.0 | Service, operations and maintenance requirements
5.1 | The supplier shall supply prices for the following in appendix 3.0 Price schedule - Service. - Attendance Price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
5.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
5.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
5.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
5.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
5.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours.. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
5.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
5.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
5.9 | The supplier shall identify all components expected to be replaced during the anticipated lifetime of the equipment and indicate the expected frequency of replacement. This includes light sources, moving parts, and other consumable or wear components. The offer shall include a complete spare parts list. | A
5.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
5.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
6.0 | Warranty requirements
6.1 | The offered warranty period shall be stated. | A
6.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.0 Price Schedule - extended warranty option | A
6.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. requirement 5.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
6.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost | A

## Subcontract 1 High shear granul
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 029.03.017 High Shear Granulator
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | General requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Inteded use: The equipment shall be intended for producing granules for pharmaceutical formulations using high-shear wet granulation of powders. It shall be suitable for use in a small-scale pharmaceutical research and teaching laboratory. | A
1.5 | The equipment shall support operation with different bowl sizes to accommodate varying batch sizes. The equipment shall do batch sizes between approximately 0.3 - 4 liters. Documentation shall be provided. | A
1.6 | The offer shall include 4 bowls suitable for different batch sizes in the range 0.3 - 4 liters. | A
1.7 | Please specify whether the high-shear mixer is configured with a top drive or a bottom drive. | I
1.8 | The equipment should allow observation of the granulation process during operation for teaching purposes, preferably through a transparent lid or window. Please describe how the granulation process can be observed. The solution will be evaluated based on the described visibility and usability for demonstration and teaching. | B2
1.9 | The granulator shall include an integrated pump for adding granulation fluid. Documentation shall be provided. | A
1.10 | The granulator shall be equipped with a spray system to uniformly add fluid to the powder during operation. Documentation shall be provided. | A
1.11 | Uniformity of fluid addition during operation should be ensured. Please provide a description of the spray system. Uniform fluid addition will be considered an advantage. The solution will be evaluated based on the uniformity of fluid addition, the distribution of fluid on the powder, and the usability of the described system. | B3
1.12 | The granulation liquid flow rate shall be adjustable. Documentation shall be provided. | A
1.13 | The equipment shall have adjustable mixing speed. Please specify the speed range. Documentation shall be provided. | A
1.14 | The granulator should be able to ensure high uniformity of granule size after granulation. Please specify the achievable granule size ranges and associated variance in millimetres. The solution will be evaluated based on the described granule size range and uniformity. | B2
1.15 | The equipment shall have a temperature sensor. Documentation shall be provided. | A
1.16 | The solution should prevent the temperature of the product from exceeding approximately 60 °C during a regular granulation procedure to prevent overheating. Please describe the proposed solutions for maintaining suitably low operation temperatures. The solution will be evaluated based on the effectiveness and usability of the described temperature control measures to conduct regular granulation. | B2
1.17 | Ports to accommodate additional probes, such as NIR or IR probes, should be available. The availability and flexibility of such ports will be considered an advantage and will result in a higher evaluation score. | B1
1.18 | The equipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.19 | The supplier shall complete the sheet Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Training requirement
2.1 | The necessary superuser training shall be provided for approximately 2-3 users. The training shall be conducted by agreement and on-site. The training should ensure safe, secure, and correct use of the equipment as well as correct maintenance to ensure longevity of the equipment. The supplier should anticipate that training will take place after handover. The scope and content of the training shall be described | A
3.0 | Hygiene/Cleaning requirements
3.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
3.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
3.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
4.0 | IT-Requirement
4.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
5.0 | Service, operations and maintenance requirements
5.1 | The supplier shall supply prices for the following in appendix 3.0 Price schedule - Service. - Attendance price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
5.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
5.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
5.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
5.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
5.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
5.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
5.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
5.9 | Describe which parts are expected to be replaced during the equipment's expected lifetime and the frequency of replacement. For example, light sources, moving parts, or other components. Complete spare part list should be included in the offer. | A
5.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
5.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
6.0 | Warranty
6.1 | The offered warranty period shall be stated. | A
6.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.0 Price Schedule - extended warranty option | A
6.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. requirement 5.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
6.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost. | A

## Subcontract 2 Rotary tablet pre
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 029.03.025 Rotary tablet press
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | General requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Inteded use: The rotary tablet press shall be used for tablet preparation and be suitable for small-scale pharmaceutical research and supervised student practicals of compression of powders and granular materials. | A
1.5 | The rotary press shall be equipped with 8 - 12 stations. Documentation shall be provided. | A
1.6 | The rotary press should have effective dust handling suitable for connection to a central dust collection system. Please describe the solution. The solution will be evaluated based on users’ assessment of usability. | B2
1.7 | The rotary press should be suitable for small batch sizes. It is considered advantageous that only a minimal number of tablets at the beginning of a batch need to be discarded. Please describe the lowest batch sizes at which uniform tablet mass is achieved. The solution will be evaluated based on the described performance. | B2
1.8 | It shall be possible to blind stations so that the press can operate with as few as one active station. The offer shall include all accessories required to blind stations to enable operation with a single active station. | A
1.9 | The rotary tablet press shall have a maximum compaction force of at least 40 kN. Documentation shall be provided. | A
1.10 | The compression process steps should be visible for teaching purposes. Clear and unobstructed visibility is considered an advantage. Please describe the proposed solution. The evaluation will be based on users’ assessment of the solution. | B2
1.11 | The tablet press shall be equipped with transparent safety guarding. | A
1.12 | The rotary press should have rapid tool-changing capabilities. Please describe the routine for changing or adding tools. The solution will be evaluated based on the described tool-changing procedure and the time required to perform it | B1
1.13 | The rotary press shall be compatible with standard EU Type B tooling. Documentation shall be provided. | A
1.14 | Compatible EU Type B tooling sets shall be supplied with the offer. Each set shall consist of flat-faced punches and a die for producing circular tablets with a diameter of 10 mm, and shall be priced in Appendix 3.1 Price Schedule - Equipment Option. | A
1.15 | The system should allow measurement of compaction force, ejection force, and displacement on individual compaction stations during production. Please describe how these measurements are performed. The solution will be evaluated based on the types of measurements possible and their accuracy. | B2
1.16 | A monitor for observing the compaction process and controlling the instrument shall be supplied. | A
1.17 | The equipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.17 | The supplier shall complete the sheet Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Training requirement
2.1 | The necessary superuser training must be provided for approximately 2-3 users. The training shall be conducted by agreement and on-site. The training should ensure safe, secure, and correct use of the equipment as well as correct maintenance to ensure longevity of the equipment. The supplier should anticipate that training will take place after handover. The scope and content of the training shall be described. | A
3.0 | Hygiene/Cleaning requirements
3.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
3.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
3.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
4.0 | IT-Requirement
4.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
5.0 | Service, operations and maintenance requirements
5.1 | The supplier shall supply prices for the following in appendix 3.1 Price schedule - Service. - Attendance price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
5.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
5.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
5.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
5.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
5.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
5.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
5.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
5.9 | Describe which parts are expected to be replaced during the equipment's expected lifetime and the frequency of replacement. For example, light sources, moving parts, or other components. Complete spare part list should be included in the offer. | A
5.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
5.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
6.0 | Warranty requirements
6.1 | The offered warranty period shall be stated. | A
6.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.1 Price Schedule - extended warranty option | A
6.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. requirement 5.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
6.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost. | A

## Subcontract 2 Single punch
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 029.03.023 Single Punch Tablet Press
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | General requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Inteded use: The single punch eccentric tablet press shall be used for tablet preparation and be suitable for small-scale pharmaceutical research and supervised student practicals of compression of powders and granular materials. | A
1.5 | The main steps of the compression process (die filling, compression, tablet ejection) shall be visibly observable. | A
1.6 | To protect operators from hazards during operation, the tablet press shall have safety guarding. The safety guarding shall be transparent. Documentation shall be provided. | A
1.7 | The tablet press shall have a powered continuous operation, with automated die filling capabilities (hopper and shoe) and tablet ejection. Documentation shall be provided. | A
1.8 | Compaction speed shall be adjustable and capable of reaching atleast 2000 tablets/hour. Documentation shall be provided. | A
1.9 | Filling depth shall be adjustable in the ca. 4 to 15 mm filling depth. Documentation shall be provided. | A
1.10 | Compression force/distance shall be adjustable and capable of reaching up to approximately 30 kN compression force. Documentation shall be provided. | A
1.11 | The single punch machine shall be operated with standard EU type B tooling. Documentation shall be provided. | A
1.12 | Compatible EU Type B tooling sets shall be supplied with the offer. Each set shall consist of flat-faced punches and a die for producing circular tablets with a diameter of 6 mm, and shall be priced in Appendix 3.1 Price Schedule - Equipment Option. | A
1.13 | Compatible EU Type B tooling sets shall be supplied with the offer. Each set shall consist of flat-faced punches and a die for producing circular tablets with a diameter of 8 mm, and shall be priced in Appendix 3.1 Price Schedule - Equipment Option. | A
1.14 | Compatible EU Type B tooling sets shall be supplied with the offer. Each set shall consist of flat-faced punches and a die for producing circular tablets with a diameter of 10 mm, and shall be priced in Appendix 3.1 Price Schedule - Equipment Option. | A
1.15 | Compatible EU Type B tooling sets shall be supplied with the offer. Each set shall consist of concave punches and a die for producing concave tablets with a diameter of 6 mm, and shall be priced in Appendix 3.1 Price Schedule - Equipment Option. | A
1.16 | The single punch machine shall be mobile and be equiped with lockable caster wheels. Documentation shall be provided. | A
1.17 | Data collection of press force and punch displacement shall be possible and exportable. The acquisition system shall be described. | A
1.18 | The machine shall have a maximum foot-print of ca. 1000x900mm. Documentation shall be provided. | A
1.19 | The equipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.20 | The supplier shall complete sheet 12 -Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Training requirement
2.1 | The necessary superuser training must be provided for approximately 2-3 users. The training shall be conducted by agreement and on-site. The training should ensure safe, secure, and correct use of the equipment as well as correct maintenance to ensure longevity of the equipment. The supplier should anticipate that training will take place after handover. The scope and content of the training shall be described. | A
3.0 | Hygiene/Cleaning requirements
3.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
3.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
3.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
4.0 | IT-Requirement
4.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
5.0 | Service, operations and maintenance requirements
5.1 | The supplier shall supply prices for the following in appendix 3.1 Price schedule - Service. - Attendance price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
5.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
5.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
5.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
5.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
5.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
5.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
5.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
5.9 | Describe which parts are expected to be replaced during the equipment's expected lifetime and the frequency of replacement. For example, light sources, moving parts, or other components. Complete spare part list should be included in the offer. | A
5.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
5.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
6.0 | Warranty requirements
6.1 | The offered warranty period shall be stated. | A
6.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.1 Price Schedule - extended warranty option. | A
6.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. requirement 5.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
6.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost. | A

## Subcontract 3 Compaction
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 029.03.024 Compaction simulator
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | General requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Inteded use: The compaction analyzer shall be used for tablet prototyping in pharmaceutical research and for supervised student practicals. | A
1.5 | The compaction analyzer shall have an automatic feeding mode that makes it possible to use it as a single-punch press for small batches, with automated repetitive compressing cycle (die filling, compression, tablet ejection). Documentation shall be provided. | A
1.6 | The compaction analyzer shall be controlled via dedicated software, allowing for defined and reproducible operation. Documentation shall be provided. | A
1.7 | It shall be possible to define, specify, and control tablet compression parameters, as well as compression profiles, through the software. Documentation shall be provided. | A
1.8 | The tablet press shall be equipped with transparent safety guarding. | A
1.9 | It shall be possible to view the main steps of the compressing cycle during processing. | A
1.10 | It shall be possible to measure and aquire data concerning the upper and lower tool displacement and force, as well as ejection force using sensors for individually compacted tablets. Documentation shall be provided. | A
1.11 | The maximum compaction force shall exceed 30 kN for both the pre-compression stage and the main compression stage. Documentation shall be provided. | A
1.12 | The compaction analyzer shall be equipped with a hopper for die filling. | A
1.13 | The compaction analyzer shall support manual powder filling. | A
1.14 | The equipment should be equipped with a forced feeder system. Please describe how forced feeding can be achieved. The solution will be evaluated based on the type of feeding system, its ability to ensure stable and reproducible sample feeding, and the operational robustness of the feeding mechanism, as documented by the supplier. User evaluation of the suitability and robustness of forced feeding will be included in the assessment. | B3
1.15 | The offer shall include a complete stationary desktop computer system with all necessary software. The computer system shall be capable of operating the equipment and acquiring, storing, and displaying measurement data for evaluation. It shall also support internet connectivity for data transfer and software updates. A display of approximately 27 inches shall be included. | A
1.16 | The software supplied with the computer system shall be capable of connecting to, acquiring, and handling data from equipment relevant to tablet quality control, including but not limited to balances, micrometers, and tensile strength testers. Documentation shall be provided. | A
1.17 | The compaction analyzer should be capable of mimicking tablet presses with different compression force profiles. This capability will be evaluated based on the range of supported force profiles and the flexibility of the control system, as described by the supplier | B2
1.18 | The compaction alalyzer shall be compatible with standard EU Type B tooling. Documentation shall be provided. | A
1.19 | The equipment should also be compatible with EU Type D tooling. It should allow easy changing between toolings. Please specify which toolings are compatible and describe the procedure for changing between them. The solution will be evaluated based on tooling flexibility, the described ease of changing tools, and the ability to accommodate multiple tooling types, including EU Type D. | B2
1.20 | The compaction simulator shall be capable of preparing mini tablets. The offer shall include two sets of multi-tip dies and punches for producing mini tablets with a 2 mm diameter, as well as any additional equipment or tooling necessary to use the multi-tip sets. Documentation shall be provided. | A
1.21 | The equipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.22 | The supplier shall complete the sheet Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Training requirement
2.1 | The necessary superuser training must be provided for approximately 2-3 users. The training shall be conducted by agreement and on-site. The training should ensure safe, secure, and correct use of the equipment as well as correct maintenance to ensure longevity of the equipment. The supplier should anticipate that training will take place after handover. The scope and content of the training shall be described. | A
3.0 | Hygiene/Cleaning requirements
3.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
3.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
3.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
4.0 | IT-Requirement
4.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
5.0 | Service, operations and maintenance requirements
5.1 | The supplier shall supply prices for the following in appendix 3.2 Price schedule - Service. - Attendance price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
5.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
5.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
5.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
5.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
5.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
5.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
5.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
5.9 | Describe which parts are expected to be replaced during the equipment's expected lifetime and the frequency of replacement. For example, light sources, moving parts, or other components. Complete spare part list should be included in the offer. | A
5.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
5.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
6.0 | Warranty requirements
6.1 | State the offered warranty period. | A
6.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.2 Price Schedule - extended warranty option | A
6.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. requirement 5.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
6.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost. | A

## Technical specifications
Techincal specifications and interfacing
Article number | Article name | Suppliers model | Acustic noise: (dB (A)) | Heat: Standby / normal operation (kW) | Room requirements: Temperature (oC peak-to-peak limits) Relative humidity min/maks (% rH) | Power supply: (schuko, fixed installation) Power usage (kW) Voltage (V) Current (A) Separate course? Separate earthing/grounding? | Network interface: Connection type Other network requirements | Pressurized gas connections: Gass (type) Pressure(bar) Purity (%) | Exhaust/ventilation: Flow (m3/hour) Diameter of connecting tube (cm) | Physical size specifications Dimensions (WxDxH)
029.03.050 | Drum Coater
029.03.017 | High Shear Granulator
029.03.025 | Rotary Tablet Press
029.03.023 | Singel Punch Excenter Machine
029.03.024 | Compaction Mimulator
029.03.007 | Fluid Bed Granulator