## Guidance for use
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Project 1105601 UiO Livsvitenskap brukerutstyr
K946.01.02 Pharmaceutical dissolution equipment
Tenderer:
HOW TO FILL OUT THE DOCUMENT
Please note that there are several tabs to be filled out.
Columns and rows are not to be deleted.
The tenderer shall only fill out fields with a white background. Fields of other colors are not to be used or altered by the Tenderer.
The tenderer shall answer YES or NO to the question on whether or not a requirement is fulfilled. All rows must be answered.
The tenderer shall answer the defined requirements. The answer shall have sufficient information for the contractor to carry out a full assessment of the tender without the need to open attachments. If the tenderer needs to provide additional or more detailed information than this document allows for, it is allowed to refer to attachments. In case of reference to attachments, the tenderer shall provide the corresponding attachment nr and relevant pages.
Minimum requirements (A-criteria) are not related to an award criterion. Lack of fulfillment of minimum requirement can lead to rejecton of the tender, cf Tilbudsinvitasjon (Tender invitation).
B-criteria are related to an award criterion. The tenderers answer to B-criteria are the basis of the contractors evaluation and assessment. B-criteria are defined as B1, B2 or B3 criteria in the corresponding column for each criterion. The definitions indicate the assigned weight of the award criterion in the tender evaluation. Award criteria indicated as B3 are assigned most weight, while the B1 criteria are assigned the least weight.
Information criteria (I-criteria) are solely for informative purposes, and are not part of the evaluation.
| Valuta: | Lisens:
| NOK | Systemlisens
| SEK | Brukerlisens
| DKK | Utstyrslisens
| EUR | Analyselisens
| USD | Annet
GBP

## Subcontract 1 Diss. fraction
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 029.03.033:1 Dissolution tester fraction
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | Quality and functionality requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Inteded use: The dissolution tester shall be used for dissolution testing in a pharmaceutical research and teaching lab. | A
1.5 | The dissolution tester shall be capable of USP type 1 & 2 configuration. All equipment needed for these configurations must be included in the offer. Documentation shall be provided. | A
1.6 | The dissolution tester shall have at least 6 vessels capable of containing 500-1000 mL. Documentation shall be provided. | A
1.7 | The dissolution tester should support retrofitting with vessels of volumes below 500 mL. Increased flexibility, including compatibility with multiple vessel sizes and straightforward retrofitting of existing systems, will be considered an advantage and scored higher. | B1
1.8 | The dissolution tester shall have a heated water bath for the vessels, with temperature control from room temperature (about 20 to 25 °C) + 5°C up to at least 37 °C. Documentation shall be provided. | A
1.9 | The heated water bath should allow temperature control from above 37 °C up to approximately 50 °C. Please specify the achievable temperature control range. Capability to reach higher temperatures will be considered advantageous and scored accordingly. | B2
1.10 | The stirring mechanism shall be electronically adjustable with a speed up to at least 150 rpm. Documentation shall be provided. | A
1.11 | The head of the dissolution tester shall be able to be raised and lowered electronically. | A
1.12 | The dissolution tester shall be operable with the head raised to a high position enabeling access to the chambers during dissolution testing. | A
1.13 | The equipment should be operable with the head lowered to a fully closed position. Please describe the solution. The ability to operate the equipment in this configuration is considered an advantage and will be evaluated based on the described functionality and operational flexibility. | B2
1.14 | The offer shall include shafts with lenghts that allow for operation both with the head in a raised position and in a fully closed position. | A
1.15 | The system shall allow sample withdrawal both manually (via cannula or shafts) and automatically, into HPLC vials and 96-well plates, without requiring system reconfiguration. Documentation shall be provided. | A
1.16 | Sample collection shall be performed using a piston-pump system. Documentation shall be provided. | A
1.17 | Filtering of samples during automatic collection should be possible. This capability is considered an advantage and will be evaluated based on the effectiveness of the filtration, compatibility with the automatic collection system, and suitability for the intended application. | B3
1.18 | Volume accuracy and repeatability of sample collection should be ensured. Please provide a full description of the automatic sample collection system, including its accuracy and pump configuration. The proposed solution will be evaluated based on sampling accuracy, repeatability, and the described pump setup." | B3
1.19 | The water bath should be easy to empty, fill, and clean. Please provide a description of the proposed solution. Ease of emptying, filling, and cleaning will be considered an advantage. The solution will be evaluated based on usability, simplicity, and the level of manual intervention required. User feedback on the solution will be included in the evaluation, with more user-friendly solutions receiving higher scores. | B1
1.20 | The equipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.21 | The supplier shall complete the sheet Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Hygiene/Cleaning requirements
2.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
2.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
2.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
3.0 | IT-Requirement
3.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
4.0 | Service, operations and maintenance requirements
4.1 | The supplier shall supply prices for the following in appendix 3.0 Price schedule - Service. - Attendance Price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
4.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
4.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
4.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
4.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
4.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
4.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
4.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
4.9 | The supplier shall identify all components expected to be replaced during the anticipated lifetime of the equipment and indicate the expected frequency of replacement. This includes light sources, moving parts, and other consumable or wear components. The offer shall include a complete spare parts lis | A
4.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
4.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
5.0 | Warranty requirements
5.1 | The offered warranty period shall be stated. | A
5.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.0 Price Schedule - extended warranty option | A
5.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. requirement 4.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
5.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost. | A

## Subcontract 2 Diss. probes
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 023.05.161 Dissolution tester (probes)
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | General requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Intended use: The dissolution tester shall be used for real-time dissolution and permeation testing of active pharmaceutical ingredients from drug formulations. | A
1.5 | The equipment shall measure concentration by using UV-VIS probes, and shall include vessels for both dissolution and permeation testing. Documentation describing the equipment and its capabilities shall be included with the offer. | A
1.6 | The equipment shall support simultaneous dissolution testing using eight vessels. Documentation shall be provided. | A
1.7 | The equipment shall support working volumes per dissolution vessel in the range of 2 mL to 20 mL. Documentation shall be provided. | A
1.8 | The equipment shall be supplied with vessels and all necessary accessories to enable permeation measurements. | A
1.9 | Permeation testing shall be performed in vessels that can accommodate the relevant permeation membranes. Documentation shall be provided. | A
1.10 | The offered solution should enable measurement of concentrations in both the donor and acceptor chambers during permeation testing. Please describe the method used to perform these measurements. Solutions providing enhanced functionality that allows concentration measurements in both donor and acceptor chambers should receive a higher evaluation score. | B3
1.11 | In-vessel UV-VIS measurement shall be observable in real time. | A
1.12 | UV–Vis measurements shall be performed independently and simultaneously in all vessels. Documentation shall be provided. | A
1.13 | UV-VIS-measurement shall provide absorbance measurements that have been automatically corrected for turbidity and other disturbances that could potentially impact the measured absorbance during the measurement. Please specify how this is corrected for in the software. Documentation shall be provided. | A
1.14 | The solution should allow the use multiple path lengths for the UV–Vis probes. Please specify the available path lengths and their approximate concentration ranges for measurement. Evaluation will be based on the number of supplied path length options, the covered concentration ranges, and the flexibility to select or chang path lengths. Solutions offering a wide range of path lengths with clear software integration and supproting documentation will receive a higher evaluation score. | B3
1.15 | Mixing shall be performed with electronically adjustable speed control. Documentation shall be provided. | A
1.16 | It shall be possible to draw samples from the vessels with cannulas during the dissolution process. | A
1.17 | Collected absorption data shall be exportable in in commonly used formats. | A
1.18 | The system shall allow temperature control in the vessels up to 40 °C. The functionality shall be described, including how temperature is controlled and maintained. | A
1.19 | The equipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.20 | The supplier shall complete the sheet Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Hygiene/Cleaning requirements
2.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
2.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
2.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
3.0 | IT-Requirement
3.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
4.0 | Service, operations and maintenance requirements
4.1 | The supplier shall supply prices for the following in appendix 3.0 Price schedule - Service. - Attendance price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
4.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
4.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
4.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
4.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
4.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
4.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
4.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
4.9 | Describe which parts are expected to be replaced during the equipment's expected lifetime and the frequency of replacement. For example, light sources, moving parts, or other components. Complete spare part list should be included in the offer. | A
4.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
4.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
5.0 | Warranty
5.1 | The offered warranty period shall be stated. | A
5.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.0 Price Schedule - extended warranty option. | A
5.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. requirement 4.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
5.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost. | A

## 12. Technical specifications
Techincal specifications and interfacing
Article number | Article name | Suppliers model | Acustic noise: (dB (A)) | Heat: Standby / normal operation (kW) | Room requirements: Temperature (oC peak-to-peak limits) Relative humidity min/maks (% rH) | Power supply: (schuko, fixed installation) Power usage (kW) Voltage (V) Current (A) Separate course? Separate earthing/grounding? | Network interface: Connection type Other network requirements | Pressurized gas connections: Gass (type) Pressure(bar) Purity (%) | Exhaust/ventilation: Flow (m3/hour) Diameter of connecting tube (cm) | Physical size specifications Dimensions (WxDxH)
029.03.050 | Film-coater for tablets
029.03.033:1 | Dissolution tester fraction
029.03.049 | Hot melt extruder-granulator
029.03.017 | Wet granulator
029.03.025 | Rotary press
029.03.024 | Compaction analyzer
029.03.007 | Fluid Bed Granulator
029.03.027 | Pellet extruder-spheronizer
029.03.023 | Singel punch exc. machine
023.05.161 | Dissolution tester (Probes)