## 1. Guidance for use
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Project 1105601 UiO Livsvitenskap brukerutstyr
K946.01.01,04 Fluid Bed Granulator
Tenderer:
HOW TO FILL OUT THE DOCUMENT
Please note that there are several tabs to be filled out.
Columns and rows are not to be deleted.
The tenderer shall only fill out fields with a white background. Fields of other colors are not to be used or altered by the Tenderer.
The tenderer shall answer YES or NO to the question on whether or not a requirement is fulfilled. All rows must be answered.
The tenderer shall answer the defined requirements. The answer shall have sufficient information for the contractor to carry out a full assessment of the tender without the need to open attachments. If the tenderer needs to provide additional or more detailed information than this document allows for, it is allowed to refer to attachments. In case of reference to attachments, the tenderer shall provide the corresponding attachment nr and relevant pages.
Minimum requirements (A-criteria) are not related to an award criterion. Lack of fulfillment of minimum requirement can lead to rejecton of the tender, cf Tilbudsinvitasjon (Tender invitation).
B-criteria are related to an award criterion. The tenderers answer to B-criteria are the basis of the contractors evaluation and assessment. B-criteria are defined as B1, B2 or B3 criteria in the corresponding column for each criterion. The definitions indicate the assigned weight of the award criterion in the tender evaluation. Award criteria indicated as B3 are assigned most weight, while the B1 criteria are assigned the least weight.
Information criteria (I-criteria) are solely for informative purposes, and are not part of the evaluation.
| Valuta: | Lisens:
| NOK | Systemlisens
| SEK | Brukerlisens
| DKK | Utstyrslisens
| EUR | Analyselisens
| USD | Annet
GBP

## 2. Fluid bed granualtor
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Article name | Art. 029.03.007 Fluid bed granulator
point nr | Description of point | Point type | Supplementary information/Additional information | Fulfillment of point Yes/No | Supplier response
1.0 | General requirements
1.1 | The instrument shall be CE-marked and shall comply with all relevant directives and standards for pharmaceutical production. A Declaration of Conformity for CE marking shall be attached to the offer | A
1.2 | The supplier shall provide an estimated delivery time from order dispatch to equipment delivery. | A
1.3 | Complete user and service manuals in English shall accompany the delivery. | A
Equipment requirements
1.4 | Inteded use: The fluid bed granulator shall be used in the process of producing and coating granules for solid dosage forms. The equipment shall be suitable for small-scale pharmaceutical research and supervised student practicals. | A
1.5 | The fluid bed shall accommodate batch sizes in the range of 100 g to 400 g. Documentation shall be provided. | A
1.6 | The fluid bed should be capable of processing batch sizes for granulation down to approximately 10 g. Suppliers shall state and document the minimum batch size that the system can reliably process. Systems capable of reliably processing smaller batch sizes will be scored higher | B3
1.7 | The fluid bed should be capable of processing batch sizes up to approximately 1000 g. Suppliers shall state and document the maximum batch size that the system can reliably process. Systems capable of reliably processing larger batch sizes will receive a higher evaluation score. | B1
1.8 | The fluid bed shall have an inlet air temperature control from room temperature (approximately 20 - 25 °C) up to approximately 80 °C. Documentation shall be provided. | A
1.9 | All parts of the fluid bed that come into contact with the product shall be made of stainless steel. | A
1.10 | Recording of inlet air temperature, exhaust air temperature, and airflow during the fluid bed process shall be possible using onboard sensors. The recorded data shall be exportable in a compatible digital format. Documentation shall be provided. | A
1.11 | The system should provide adjustable inlet air flow up to at least 60 m³/h. Please provide documentation for the following parameters: the maximum achievable airflow (m³/h), control accuracy (± %), and a description of the control method. All values shall be supported by datasheets, test reports, or equivalent documentation. The evaluation will be based on the described range of achievable air flow up to at least 60 m³/h and accuracy of this control. Solutions providing wider range and more accurate control will achieve higher scores. | B2
1.12 | For the inlet air, please specify the suitable range of humidities of the inlet air in percentage relative humidity and describe whether it is possible to draw the inlet air from the room the equipment is operated in. | I
1.13 | For the inlet air, please specify whether a filter is connected to the inlet air draw such as to ensure the necessary low-particle quality of the inlet air. If no such filter is connected, please describe the suitable particulate ranges for the inlet air to be supplied to the equipment during operation. | I
1.14 | For the outlet air, please specify the maximum necessary air to be exhausted during operation in Liters per minute. This will be used to access whether the equipment needs to be connected to a ventilation channel for exhaust. In addition, please describe how the system can be connected to an exhaust outlet and what type of connector to be used. | I
1.15 | The fluid bed granulator shall be designed to allow connection to a ventilation system for the extraction of toxic or harmful fumes. Documentation of the solution shall be included with the tender. | A
1.16 | Product retention filters and process filters should be suitable for processing pharmaceuticals. Solutions using robust, corrosion-resistant, pharmaceutical-grade materials (e.g. documented stainless steel grades suitable for GMP use) will receive a higher evaluation score. Please describe and specify the materials used for the product retention and process filters, including material grade and surface finish where applicable. | B2
1.17 | The fluid bed system shall include integrated methods to address filter clogging, such as back-pulsing, mechanical shaking of the filters, or other equally effective solutions. The method(s) shall be described, together with an explanation of how it ensures reliable system operation. | A
1.18 | The liquid for the nozzle shall be fed with an adjustable flow rate by a peristaltic pump. Documentation shall be provided. | A
1.19 | The spray liquid used in the nozzle will consist of aqueous and ethanol-based solutions. The system shall be compatible with both types of solutions. The supplier shall provide documentation demonstrating compatibility with these liquids. | A
1.20 | A wide range of compatible nozzles and exchangeable nozzle tips should be provided to support flexible operation of the fluid bed system. Please describe the available nozzle types and exchangeable tips, including their compatibility, intended applications, and any limitations. Systems offering a broad selection of nozzle types, ease of interchangeability, and suitability for diverse process applications should receive higher evaluation scores. | B1
1.21 | Prices for compatible nozzles and exchangeable nozzle tips shall be listed in Appendix 3.0 Price Schedule – Add-ons/Equipment, in accordance with Requirement 1.20 | A
1.22 | The atomizing air of the nozzle shall be adjustable. Please describe how this is achieved and specify which parameters can be adjusted. Documentation shall be provided. | A
1.23 | The fluid bed shall be capable of operating with both top and bottom spray configurations, as well as in Wurster (bottom-spray coating) configuration. Documentation shall be provided. | A
1.24 | The equipment must be compatible with a 230V single-phase or 400V three-phase electrical supply. | A
1.25 | The supplier shall complete the sheet Technical specifications for this article, providing all relevant technical data and parameter values. | A
2.0 | Training requirement
2.1 | The necessary superuser training must be provided for approximately 2-3 users. The training shall be conducted by agreement and on-site. The training should ensure safe, secure, and correct use of the equipment as well as correct maintenance to ensure longevity of the equipment. The supplier should anticipate that training will take place after handover. The scope and content of the training shall be described. | A
3.0 | Hygiene/Cleaning requirements
3.1 | Any disinfectants or cleaning methods that should not be used on the equipment should be described, along with an explanation of why they are unsuitable. | I
3.2 | The recommended procedure(s) for cleaning, disinfection, and/or sterilization of the offered equipment shall be described. | A
3.3 | The offered equipment should be designed to facilitate easy cleaning and disinfection, preferably with as few cracks, joints, and recesses as possible. Please provide a detailed description of the cleaning and disinfection procedures, including the total time required for complete cleaning. User evaluation of the time and effort required will be included in the assessment. | B1
4.0 | IT-Requirement
4.1 | All mandatory program and security updates for both software and equipment shall be provided and installed at no cost until end of service. This requirement is only relevant if there are software associated with the equipment. Updates refer to improvements and bug fixes for the software. | A
5.0 | Service, operations and maintenance requirements
5.1 | The supplier shall supply prices for the following in appendix 3.0 Price schedule - Service. - Attendance price including all expenses and time spent per round trip to/from the service location. - Hourly Rate for Service Execution: covers the time spent at the client or supplier location. | A
5.2 | A short delivery time for spare parts is an advantage. The supplier should specify the maximum delivery time from order to delivery at the client’s location, including delivery from an external warehouse (1 calendar day = 24 hours). Shorter delivery times will be evaluated more favorably. Please describe the location of the spare parts warehouse. | B1
5.3 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual should be in English or a Scandinavian language. The client shall receive updates to the service manual, unsolicited and free of charge, until the end of service as specified by the manufacturer. | A
5.4 | Free customer support by phone shall be offered during regular business hours. Qualified technical personnel must be available to answer questions. Communication with customer support shall be conducted in a Scandinavian language or English. Regular business hours are defined as weekdays from 08:00 to 16:00. Free customer support shall be offered until the equipment reaches end of service. The type of customer support offered shall be specified. | A
5.5 | A short response time is an advantage. Specify the response time for qualified technical personnel, by phone or similar, when technical assistance is needed during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. It is not permitted to make reservations about establishing remote access to meet the response time. The response time should be specified in minutes or, if necessary, hours, and is defined as the time from when the client requests technical assistance until qualified technical personnel make contact to begin troubleshooting by phone or similar. | B1
5.6 | The supplier’s service organization should have a quick response time for urgent issues. Please specify the maximum attendance time (in hours) required for emergency service after an error report is received during regular business hours. Regular business hours are defined as weekdays from 08:00 to 16:00. | B1
5.7 | Provide a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Protocols for conducting preventive maintenance, along with information on frequency and expected time usage, shall be attached. | A
5.8 | If equipment-specific special tools, calibration tools, software, passwords, service keys, or similar items are required to perform technical troubleshooting, maintenance, and/or emergency repairs, these shall be included in the offer. | A
5.9 | Describe which parts are expected to be replaced during the equipment's expected lifetime and the frequency of replacement. For example, light sources, moving parts, or other components. Complete spare part list should be included in the offer. | A
5.10 | The supplier shall guarantee access to spare parts for at least 7 years after the delivery of the equipment. | A
5.11 | Access to spare parts beyond 7 years is emphasized. Specify the total number of years the supplier guarantees access to spare parts. | B1
6.0 | Warranty requirements
6.1 | The offered warranty period shall be stated. | A
6.2 | An extended warranty period shall be offered as an option and shall be priced in Appendix 3.0 Price Schedule - extended warranty option | A
6.3 | During the warranty period, the supplier shall provide services equivalent to Maintenance Agreement Level 1 – Preventive. Maintenance Agreement Level 1 – Preventive includes: - Preventive maintenance in accordance with the manufacturer’s specifications, including labor, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - Response-time guarantee, cf. Requirement 5.5. Documentation of completed work (service reports) shall be submitted to the contracting authority within five working days. Unless otherwise agreed, maintenance shall be performed at the user’s location. The supplier shall coordinate the timing of service delivery with the contracting authority in advance. The final periodic maintenance shall be carried out towards the end of the warranty period. For example, for a one-year warranty, one periodic maintenance shall be performed if the manufacturer recommends annual maintenance. | A
6.4 | At the end of the warranty period, the supplied equipment shall have the latest version of the software at no additional cost. | A

## 3. Technical specifications
Suppliers model | Acustic noise: (dB (A)) | Heat: Standby / normal operation (kW) | Room requirements: Temperature (oC peak-to-peak limits) Relative humidity min/maks (% rH) | Power supply: (schuko, fixed installation) Power usage (kW) Voltage (V) Current (A) Separate course? Separate earthing/grounding? | Network interface: Connection type Other network requirements | Pressurized gas connections: Gass (type) Pressure(bar) Purity (%) | Exhaust/ventilation: Flow (m3/hour) Diameter of connecting tube (cm) | Physical size specifications Dimensions (WxDxH)