Appendix 1 Description of the deliverables/requirement specification

The Buyer shall complete Appendix 1

## [The Appendix must be composed in such manner that it is clear to the Supplier what is purchased and how they shall respond to the requirement specification in their solution specification (Appendix 2). It is important that the Buyer is explicit in his information share, both in the overall description and in the table of requirements, and that the essential information for the Supplier’s solution specification is well described]

This tender concerns the acquisition of an advanced, high-throughput confocal microscopy system for the Molecular Medicine Imaging Platform (MIP) UiO Core Facility at the Institute of Basic Medical Sciences (IMB), Faculty of Medicine. The system will serve a broad and diverse user community and must be fully suited to operate in a core facility environment. Our aim is to expand our microscopy portfolio with a state-of-the-art imaging platform capable of high-throughput long-term confocal imaging of large 3D specimens as well as cell monolayers. Key evaluation criteria will include image quality, imaging speed, field of view as well as robust and stable environmental control enabling long-term live imaging. The system must also be user-friendly and capable of supporting both routine, easy protocols and advanced imaging applications.

CRITERIA INTERPRETATION

A: Obligatory requirements that must be met. Deviation from these requirements will result in evaluation of the deviation and can result in rejection of the offer. A*: A* requirement (= SHALL requirement with additional value). A* are the requirements that SHALL be fulfilled. Tenders that do not fulfil the minimum requirements risk rejection. Answers will be of great importance for the evaluation of the tender. B: Desirable options and key criteria for evaluation and ranking. The evaluation criteria and specifications will be integrated to produce a final ranking of offers.

| No. | Priority | Description |
| --- | --- | --- |
| - DESCRIPTION OF THE SYSTEM |  |  |
|  | A | The system shall be able to perform high-throughput, long-term, widefield and confocal imaging of cell monolayers as well as of larger 3D specimens such as spheroids, organoids, gastruloids and embryos (diameter up to several 100 µm). |
| Suppliers confirmation and description; |  |  |
|  | A | The system shall support imaging of complex 3D specimens over multiple days with stable performance (mechanical stability, focus stability and environmental stability). Describe any limitations regarding maximum experiment duration and performance drift over time. |
| Suppliers confirmation and description; |  |  |
|  | A | The system shall support high-throughput imaging of multi-well plates (including 96- and 384-well formats), including precise plate-based coordinate handling and plate-map definitions. Describe supported formats and mechanical constraints. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The imaging solution shall enable stable long-term climate control (multiple days), including temperature, humidity and CO₂, via a full enclosure (cage incubator or equivalent). Describe climate-control solutions and achievable stability/uniformity. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | Experimental reproducibility is critical. The system shall include design and routines that ensure experimental comparability and reproducibility. Describe the solutions (hardware calibration, environmental control, protocol transfer, configuration management). |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The system shall include integrated solutions for automated and user-friendly imaging setup. Outline the levels of user customization available for imaging setups and assays, and how non-expert users can configure standard experiments. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Reproducibility between experiments. Describe how configurations can be shared or cloned and how reproducible performance is achieved. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | The system should include customizable and easily implemented analysis pipelines within the software for common assays (e.g. organoid tracking, intensity quantification, nuclei segmentation and counting). Describe included pipelines and options for user-level customization. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | The system should be easy to operate with low entry knowledge level for both hardware and software. Describe the user interface design, available training materials, and typical time needed to train a new user (both for routine and advanced protocols). |
| Suppliers confirmation and description; |  |  |
| - IMAGING REQUIREMENTS |  |  |
| - 2 | A* | The imaging solution shall include at least 3 laser lines, including 488 nm and 561 nm (or similar), and a third line in the far-red (<700 nm). Inclusion of additional laser lines (most importantly, UV/405 nm, but also NIR) will be viewed favourably. Describe solutions and upgrade possibilities. Include nominal power and expected lifetime. |
| Suppliers confirmation and description; |  |  |
| - 2 | A | The system shall provide integrated transmitted-light acquisition. |
| Suppliers confirmation and description; |  |  |
| - 2 | A* | It shall be possible to perform at least 3-colour imaging within a single experiment acquisition protocol (sequential or simultaneous as appropriate). Describe how multi-colour imaging is implemented and how crosstalk between channels is minimized. |
| Suppliers confirmation and description; |  |  |
|  | A* | The imaging solution shall be able to image both 2D specimens (e.g. monolayer cell cultures) and 3D specimens (e.g. organoids) in high-throughput imaging setups. Describe any constraints for 2D vs 3D imaging (e.g. working distance, speed, depth limits). |
| Suppliers confirmation and description; |  |  |
| - 2 | A* | The system shall include both confocal and widefield imaging solutions. Describe the confocal principle and widefield implementation, including confocal quality, mechanism of z-stacking and minimal z-section thickness. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The system should be able to perform multi-well point revisiting to enable long-term tracking of cells and 3D structures. Describe reproducibility and accuracy in XY-point revisiting and z-focusing. Indicate minimal reproducible z-step size and maximal total z-stack depth. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The system should allow flexible control over imaging parameters such as stack height, step size, illumination intensity, exposure time, imaging frequency (time-lapse). Describe the flexibility and any limitations. |
| Suppliers confirmation and description; |  |  |
| - 2 | A* | The system must include appropriate objectives enabling imaging through both glass and plastic at low and medium magnifications (up to at least 40x). One or more high-magnification immersion objectives (60x, e.g. oil- or water-immersion, with dispenser) suitable for long-term 3D imaging are desirable. Describe the included objective set (magnification, NA, working distance, immersion type). |
| Suppliers confirmation and description; |  |  |
| - 2 | B | A large field of view (FOV) is highly desirable to reduce or avoid stitching. Ideally, a single organoid (approximate diameter 250 µm) should fit within one FOV at 20x and 40x. Specify the FOV (µm and pixels) for relevant objectives and describe stitching options. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The system should allow fast imaging of full 96-well plates in confocal mode. Describe the time required to image a full 96-well plate with 1, 2 or 3 laser lines, using representative exposure times. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The confocal solution should provide well-resolved sub-cellular information in 3D samples (e.g. spinning disk or similar). Good optical sectioning in organoids and similar models is required. Describe expected resolution (X, Y, Z) and depth performance. |
| Suppliers confirmation and description; |  |  |
|  | B | Inclusion of additional imaging hardware solutions that improve image quality and/or resolution will be viewed very favourably. |
| Suppliers confirmation and description; |  |  |
| - B | B | The availability of advanced or smart acquisition (adaptive microscopy) strategies will be considered highly favourable. Please describe the available solutions, their ease of implementation, and the realistic training time required for users to learn how to use these options. |
| Suppliers confirmation and description; |  |  |
| - DETECTORS AND DATA QUALITY |  |  |
|  | A | The system shall be equipped with detectors suitable for both transmitted-light and fluorescence imaging. |
| Suppliers confirmation and description; |  |  |
|  | A* | The system shall be equipped with at least one high-sensitivity, fast detection system with a large FOV, suitable for high-throughput imaging. The detectors shall support image acquisition with a bit depth to 16 bits. |
| Suppliers confirmation and description; |  |  |
| - 3 | B | Included camera(s) should have high quantum efficiency (≈ 90 % or higher in 400–750 nm range), low read noise (≈ 1 e⁻ or lower) and high dynamic range. Describe the offered detector(s). |
| Suppliers confirmation and description; |  |  |
| - 3 | B | The camera pixel size should be well matched to the offered high-resolution objectives to allow proper Nyquist sampling. Specify effective pixel size (µm) at relevant magnifications. |
| Suppliers confirmation and description; |  |  |
| - 3 | A | The camera shall allow fast acquisition speeds (> 50 fps at full FOV). Describe achievable frame rates at relevant frame sizes. |
| Suppliers confirmation and description; |  |  |
| - 3 | B | The camera chip should have at least 2048 × 2048 pixels. State sensor size and FOV (in µm) with the offered objectives. |
| Suppliers confirmation and description; |  |  |
|  | B | The camera should use a high-speed data interface suitable for fast image acquisition and transfer. Describe the camera connection type(s) and their maximum supported data rates. |
| Suppliers confirmation and description; |  |  |
| - CONFOCAL MODULE AND LASER SOURCES |  |  |
| - 4 | A* | The system shall be equipped with a fast confocal module providing optical sectioning of samples. Describe the confocal technology (e.g. spinning disk) and its suitability for high-throughput 3D imaging. |
| Suppliers confirmation and description; |  |  |
| - B | B | Describe the process and ease for facility staff to exchange the disks in case several spinning disks are available. |
| Suppliers confirmation and description; |  |  |
| - 4 | A | The detection system should allow to image a large FOV. Specify the maximum FOV that can be scanned with representative objectives. |
| Suppliers confirmation and description; |  |  |
| - 4 | A | The camera should support fast imaging with a large field of view. Specify confocal imaging speeds for full and reduced frame sizes (e.g. 2048 × 2048, 1024 × 1024, 512 × 512). |
| Suppliers confirmation and description; |  |  |
| - 4 | A* | The confocal module should provide deep imaging (> 100 µm where relevant) with good rejection of out-of-focus light and high confocality. Describe achievable resolution (X, Y, Z) and section thickness with the offered objectives. |
| Suppliers confirmation and description; |  |  |
| - 4 | B | The confocal module should be compatible with fast sequential multichannel acquisition. Describe relevant features. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe features that minimize photo-damage to the sample (e.g. laser modulation, adaptive illumination). |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the solution for separation of laser excitation and fluorescence emission light (e.g. dichroics, emission filters). |
| Suppliers confirmation and description; |  |  |
| - 4 | B | All lasers should be solid-state or similar technologies, with long lifetime, low maintenance and sufficient power for fast imaging. Describe expected durability and minimum guaranteed power output per laser. |
| Suppliers confirmation and description; |  |  |
| - 4 | B | The system should be upgradable with additional lasers; if field-upgrades are not possible, provide options for alternative laser sources to cover additional fluorophores (e.g. CFP, YFP, infrared dyes). |
| Suppliers confirmation and description; |  |  |
| - OBJECTIVE CONSTELLATION OPTIONS |  |  |
|  | A* | The objective set must be suitable for imaging live and fixed cells, tissues, organoids (150–400 µm depth) and small organisms (e.g. Drosophila, zebrafish). The offer must include at least one high-quality high-magnification objective (e.g. 40x–60x) suitable for 3D imaging of organoids and similar samples. |
| Suppliers confirmation and description; |  |  |
|  | A | List, detail and price objective options that are well suited for the models described above. Options should include low-magnification (≈ 4x) and higher-magnification objectives (≈ 40x–100x). If possible, include high-end water or silicone-immersion objectives (minimum working distance 300 µm) or suitable immersion lenses; specify effective NA in different immersion media. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the maximum number of objectives that can be mounted in the turret simultaneously. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Detail DIC equipment for the offered objective lenses where relevant. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Objectives should have a high degree of chromatic aberration correction suitable for confocal imaging. Describe the wavelength range over which the objectives are corrected. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Objectives should be corrected (flat field and chromatic aberration) over a large FOV. Specify the corrected field size. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Objective constellation should include at least one objective with long working distance and one high-quality high numerical aperture objective; Specify for each objective. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Provide a list of additional dry, water, oil and silicone-immersion objectives suitable for the system and anticipated user needs. |
| Suppliers confirmation and description; |  |  |
| - B | B | Describe the ease of use and maintenance of different objectives, particularly in cases where user interaction is required (e.g. correction collar adjustment, use of a water dispenser). |
| Suppliers confirmation and description; |  |  |
| - SAMPLE MOUNTING AND CLIMATE CONTROL |  |  |
| - 6 | A | The system shall be equipped with a heated cage incubator (or similar) providing stable environmental control suitable for extended live-cell imaging experiments. |
| Suppliers confirmation and description; |  |  |
| - 6 | A | The system shall include the necessary equipment to control temperature, CO₂ and humidity. |
| Suppliers confirmation and description; |  |  |
| - 6 | A | Environmental control units shall provide accurate and active temperature control, digital/active CO₂/air mixing and active humidity control. Objective heaters should be available if deemed necessary. Describe the proposed solutions. |
| Suppliers confirmation and description; |  |  |
| - 6 | A | The CO₂ controller shall be an active mixing unit capable of mixing 100 % CO₂ with ambient air to the desired CO₂ concentration. A simple flow-control unit is not sufficient. |
| Suppliers confirmation and description; |  |  |
|  | B | Additional environmental control capabilities (e.g. hypoxia) will be considered favourably. |
| Suppliers confirmation and description; |  |  |
| - B | B | Describe the maintenance requirements of the environmental control system, if any. |
| Suppliers confirmation and description; |  |  |
| - 6 | B | The heated system should include sample holders compatible with several sample formats (e.g. multi-well plates, smaller multi-well imaging chambers, individual 35 mm dishes). The system shall include a slide holder accommodating conventional glass slides (76 mm × 26 mm). Flexibility to accommodate non-conventional sample format dimensions (like custom-made chips) will be viewed positively. |
| Suppliers confirmation and description; |  |  |
| - HIGH-THROUGHPUT AND PLATE HANDLING |  |  |
| - 7 | A | The system shall support high-throughput imaging of multi-well plates (e.g. 96-wells), with precise plate-based coordinate handling and plate-map definitions. Describe supported formats and mechanical constraints. |
| Suppliers confirmation and description; |  |  |
| - 7 | B | The stage or plate-handling solution should provide high accuracy and repeatability suitable for long-term point revisiting in all wells of multi-well plates. Describe positioning accuracy and repeatability specifications, including performance when the plate is removed from the instrument (e.g. for treatment) and subsequently reinserted. |
| Suppliers confirmation and description; |  |  |
| - 7 | B | Describe any high-content/high-throughput analysis workflows integrated in the system software that are particularly suited for plate-based experiments. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe compatible plate geometries for high-throughput (e.g. round or flat bottom, low or high profile plates). |
| Suppliers confirmation and description; |  |  |
| - B | B | The system should include flexible autofocus solutions suitable for high-throughput, plate-based imaging. Describe available autofocus methods, their impact on speed, photobleaching, and robustness, as well as their compatibility with different vessel types. |
| Suppliers confirmation and description; |  |  |
| - SOFTWARE, AUTOMATION AND ANALYSIS |  |  |
| - 8 | A* | The acquisition workstation shall come with dedicated software for multi-dimensional image acquisition and processing. The software should integrate all relevant hardware and support fast, hardware-synchronized imaging operations. |
| Suppliers confirmation and description; |  |  |
|  | A | Software controlling all parts of the system (microscope, stages, climate control, illumination, detectors, etc.) shall be included. |
| Suppliers confirmation and description; |  |  |
| - 8 | B | The software should allow automated multi-dimensional experiments (channels, z, time) at multiple positions and for different plate formats, both for widefield and confocal imaging. |
| Suppliers confirmation and description; |  |  |
| - 8 | A | The software should allow for fast setup of experiments. Describe how experiment setup is supported (e.g. templates, experiment editors, plate layout tools). |
| Suppliers confirmation and description; |  |  |
| - 8 | A* | The software shall include user-friendly setup dialogs (e.g. wizards) for multi-dimensional image acquisition, stitching and multi-position imaging. Please describe solutions, their ease of implementation, and the realistic training time required for users to learn how to use these options. |
| Suppliers confirmation and description; |  |  |
| - 8 | A | The software shall be capable of supporting high-content screening experiments. Describe this functionality and how such experiments are configured. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the available analysis modules; the level of analysis they provide (e.g. mask images, CSV tables, plots, reports); the types of image data they can handle (e.g. individual planes, Z-stacks, projections, time series, 3D renderings); and whether quantitative values can be directly visualized and contrasted with the corresponding images. |
| Suppliers confirmation and description; |  |  |
| - A | B | Describe the file format(s) and folder structure used by the system, and how they are adapted to high-throughput imaging workflows. |
| Suppliers confirmation and description; |  |  |
| - A | B | Describe how metadata are stored and handled (e.g. experimental and imaging settings, plate maps, analysis results) and how easily they can be searched, retrieved, and exported for downstream use. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the solution for plate visualization (e.g. synchronized lookup tables across wells, consistent min/max intensity settings, time-lapse data visualization, overview heatmaps, and other tools that facilitate comparison of images and signals across the plate). |
| Suppliers confirmation and description; |  |  |
| - 8 | B | The software should allow smart microscopy (i.e. altering imaging parameters based on feedback from automated image analysis during acquisition). Describe the available possibilities and user configuration options, as well as their ease of implementation and the realistic training time required for users to learn how to use these options. |
| Suppliers confirmation and description; |  |  |
| - 8 | B | The system should include customizable and easily implemented analysis pipelines for common assays relevant to the described models. Describe included analysis modules and options for user-level customization. |
| Suppliers confirmation and description; |  |  |
|  | B | The system should include analysis options that make use of AI or machine learning (e.g. for segmentation, classification, object detection, or anomaly detection). Describe these AI-based analysis functionalities and how they can be configured and used by end users. |
| Suppliers confirmation and description; |  |  |
|  | B | The inclusion of additional licences for analysis workstation(s) will be viewed highly favourably. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe any solutions for importing images from third-party instruments for analysis, if available. |
| Suppliers confirmation and description; |  |  |
| - OPTIONS AND UPGRADES |  |  |
| - 9 | B | To anticipate and accommodate developing user needs, modularity and future upgradability with well-integrated modules will be viewed highly favourable. Include a list of current and expected (2-year horizon) hardware and software upgrades, how they integrate with the offered system and the complexity of installation (duration, on-site vs factory, compatibility). |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Alternative and additional cameras suitable for the system, including options with smaller pixel size. Provide options and prices. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Differential interference contrast (DIC) solutions. Describe possible DIC configurations and associated components. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Fast z-scanning solutions such as piezo stages. Describe available options, integration and pricing. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Warranty extension. Give prices per year and indicate for how long the warranty can be extended. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Preventive maintenance contract (not covering parts). Give prices per year and indicate for how long a maintenance contract is available. |
| Suppliers confirmation and description; |  |  |
| - WORKSTATION AND SOFTWARE |  |  |
|  | A | The offer should include a high-end workstation suitable for the system, with appropriate memory, fast temporary storage and long-term storage for large time-lapse datasets and high-frame-rate imaging. Detail the minimal specifications required to guarantee full system functionality. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The offer should include one large or two medium-sized monitors that are ergonomic and have high resolution. Detail the minimal specifications required of such screens. |
| Suppliers confirmation and description; |  |  |
|  | B | The acquisition workstation should come with the software described in Section 8 and must be configured to support hardware-synchronized multi-dimensional imaging. |
| Suppliers confirmation and description; |  |  |
| - ANTI-VIBRATION TABLE |  |  |
| - 1 | B | Describe options and prices for active anti-vibrational table including silent compressor if one is required for the instrument. |
|  | B | Future upgrades to the system (i.e. additional optical components or modules) should fit on the anti-vibration table. Please state available size options (dimensions, load capacity, configurability) and dimensions required to incorporate future large hardware extensions. |
| Suppliers confirmation and description; |  |  |
| - DELIVERY AND SERVICE |  |  |
| - 1 | A | Delivery should not exceed 6 months after received written order. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Indicate expected delivery time for the equipment as well as time requirements for installation until the equipment is ready for use. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | UiO has limited space for long-term storage and no specialized equipment for installation or internal transport. Specifically indicate any such needs. |
| Suppliers confirmation and description; |  |  |
|  | B | Provide realistic and maximum estimates for delivery time for spare parts, and specify. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Provide and specify realistic and maximum timelines for repair routines, from reception of a request from the UiO facility to on-site repair. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the contact line for technical support (phone, e-mail, general portal, etc.) and expected response times. This service should be free of charge for hardware and software throughout their lifetime. |
| Suppliers confirmation and description; |  |  |
|  | A | Include a list with prices for parts required for preventive maintenance visits, as well as spare parts. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The provider will guarantee access to user support, technical support, spare parts and full repair for at least 7 years. Indicated guarantees beyond 7 years will be evaluated positively. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Describe, in short, the maintenance routines (estimated duration and frequency) and equipment needed to be followed by the facility manager as part of local preventive maintenance routines. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe nature, content, frequency and duration of the preventive maintenance visits scheduled during the warranty period. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | All software updates needed to run the system should always be available and free of charge throughout the lifetime of the software and system. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Describe the organization, size and availability of user and technical support. Indicate the location of the closest technical support facility. Presence of local expert support will be evaluated positively. |
| Suppliers confirmation and description; |  |  |
| - WARRANTY |  |  |
|  | A | Describe the warranty solutions included in the offer and alternatives if any. Indicate the guaranteed running time (in % of working days) during the warranty period. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | Warranty must comply with Norwegian regulations and cover costs for all repairs, hardware and software needed during this process. It should also cover all costs for travel and working fees /during these visits. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | During the warranty period, preventive maintenance visits (PMVs) must be included and without added charges. The last PMV will be near the end of the warranty and include a software update. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | List any parts that are covered by a longer warranty. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Provide an overview and pricing of service/support solutions after expiration of the warranty. |
| Suppliers confirmation and description; |  |  |
| - TRAINING |  |  |
|  | A | Our facility serves a large community with varying user expertise. At least 2 days of on-site user training for at least 4 users must be included. |
| Suppliers confirmation and description; |  |  |
|  | A | Additional training sessions for advanced users (up to 4 users) must be included. Describe the content and duration of such sessions. |
| Suppliers confirmation and description; |  |  |
| - ENVIRONMENTAL FOOTPRINT |  |  |
| - 1 | B | Indicate power consumption (in kWh) at normal use (10h/day). |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Indicate the extent to which equipment components can be recycled upon disposal (in %). |
| Suppliers confirmation and description; |  |  |
| 15.3 | B | Indicate the expected normal and maximum allowed noise level (in dB) of the entire system (including climate-control setup) during normal use. |
| Suppliers confirmation and description; |  |  |
| 15.4 | B | Solutions that incorporate/recycle existing surplus MIP equipment will be reviewed positively. These elements include objectives (Nikon MRD01691, MRD01905, MRH01401, MRH48430; Zeiss 420782-9900-799, 421787-9970-799, 421767-9970-000; Olympus 1-U2B836, 1-UB827, 1-U2C377), antivibration tables (120x85 or 200x150, both with metric threads), or microscope stands (Nikon Ti2-Eclipse). Describe compatibility, mode and extent of incorporation, and limitations. If needed, further details regarding equipment can be provided upon request. |
| Suppliers confirmation and description; |  |  |
| 15.5 | B | The provider will guarantee access to user support, technical support, spare parts and full repair for at least lifetime. Indicated guarantees beyond lifetime years will be evaluated positively. |
| Suppliers confirmation and description; |  |  |
| 15.6 | A* | State the guaranteed life-time period, at least 7 years, of the system which include available spare parts and software upgrades |
| Suppliers confirmation and description; |  |  |
| - ADDITIONAL CONSIDERATIONS |  |  |
|  | A | All equipment must meet the relevant directives and standards. Include separate documentation that supports this (Declaration of Conformity). |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Describe the room requirements for installation and normal operation of the equipment (consider vibrations, power outlets, temperature control, running water and gas). |
| Suppliers confirmation and description; |  |  |

Appendix 2 Supplier’s solution specification

This tender concerns the acquisition of an advanced, high-throughput confocal microscopy system for the Molecular Medicine Imaging Platform (MIP) UiO Core Facility at the Institute of Basic Medical Sciences (IMB), Faculty of Medicine. The system will serve a broad and diverse user community and must be fully suited to operate in a core facility environment. Our aim is to expand our microscopy portfolio with a state-of-the-art imaging platform capable of high-throughput long-term confocal imaging of large 3D specimens as well as cell monolayers. Key evaluation criteria will include image quality, imaging speed, field of view as well as robust and stable environmental control enabling long-term live imaging. The system must also be user-friendly and capable of supporting both routine, easy protocols and advanced imaging applications.

CRITERIA INTERPRETATION

A: Obligatory requirements that must be met. Deviation from these requirements will result in evaluation of the deviation and can result in rejection of the offer. A*: A* requirement (= SHALL requirement with additional value). A* are the requirements that SHALL be fulfilled. Tenders that do not fulfil the minimum requirements risk rejection. Answers will be of great importance for the evaluation of the tender. B: Desirable options and key criteria for evaluation and ranking. The evaluation criteria and specifications will be integrated to produce a final ranking of offers.

| No. | Priority | Description |
| --- | --- | --- |
| - DESCRIPTION OF THE SYSTEM |  |  |
|  | A | The system shall be able to perform high-throughput, long-term, widefield and confocal imaging of cell monolayers as well as of larger 3D specimens such as spheroids, organoids, gastruloids and embryos (diameter up to several 100 µm). |
| Suppliers confirmation and description; |  |  |
|  | A | The system shall support imaging of complex 3D specimens over multiple days with stable performance (mechanical stability, focus stability and environmental stability). Describe any limitations regarding maximum experiment duration and performance drift over time. |
| Suppliers confirmation and description; |  |  |
|  | A | The system shall support high-throughput imaging of multi-well plates (including 96- and 384-well formats), including precise plate-based coordinate handling and plate-map definitions. Describe supported formats and mechanical constraints. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The imaging solution shall enable stable long-term climate control (multiple days), including temperature, humidity and CO₂, via a full enclosure (cage incubator or equivalent). Describe climate-control solutions and achievable stability/uniformity. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | Experimental reproducibility is critical. The system shall include design and routines that ensure experimental comparability and reproducibility. Describe the solutions (hardware calibration, environmental control, protocol transfer, configuration management). |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The system shall include integrated solutions for automated and user-friendly imaging setup. Outline the levels of user customization available for imaging setups and assays, and how non-expert users can configure standard experiments. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Reproducibility between experiments. Describe how configurations can be shared or cloned and how reproducible performance is achieved. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | The system should include customizable and easily implemented analysis pipelines within the software for common assays (e.g. organoid tracking, intensity quantification, nuclei segmentation and counting). Describe included pipelines and options for user-level customization. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | The system should be easy to operate with low entry knowledge level for both hardware and software. Describe the user interface design, available training materials, and typical time needed to train a new user (both for routine and advanced protocols). |
| Suppliers confirmation and description; |  |  |
| - IMAGING REQUIREMENTS |  |  |
| - 2 | A* | The imaging solution shall include at least 3 laser lines, including 488 nm and 561 nm (or similar), and a third line in the far-red (<700 nm). Inclusion of additional laser lines (most importantly, UV/405 nm, but also NIR) will be viewed favourably. Describe solutions and upgrade possibilities. Include nominal power and expected lifetime. |
| Suppliers confirmation and description; |  |  |
| - 2 | A | The system shall provide integrated transmitted-light acquisition. |
| Suppliers confirmation and description; |  |  |
| - 2 | A* | It shall be possible to perform at least 3-colour imaging within a single experiment acquisition protocol (sequential or simultaneous as appropriate). Describe how multi-colour imaging is implemented and how crosstalk between channels is minimized. |
| Suppliers confirmation and description; |  |  |
|  | A* | The imaging solution shall be able to image both 2D specimens (e.g. monolayer cell cultures) and 3D specimens (e.g. organoids) in high-throughput imaging setups. Describe any constraints for 2D vs 3D imaging (e.g. working distance, speed, depth limits). |
| Suppliers confirmation and description; |  |  |
| - 2 | A* | The system shall include both confocal and widefield imaging solutions. Describe the confocal principle and widefield implementation, including confocal quality, mechanism of z-stacking and minimal z-section thickness. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The system should be able to perform multi-well point revisiting to enable long-term tracking of cells and 3D structures. Describe reproducibility and accuracy in XY-point revisiting and z-focusing. Indicate minimal reproducible z-step size and maximal total z-stack depth. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The system should allow flexible control over imaging parameters such as stack height, step size, illumination intensity, exposure time, imaging frequency (time-lapse). Describe the flexibility and any limitations. |
| Suppliers confirmation and description; |  |  |
| - 2 | A* | The system must include appropriate objectives enabling imaging through both glass and plastic at low and medium magnifications (up to at least 40x). One or more high-magnification immersion objectives (60x, e.g. oil- or water-immersion, with dispenser) suitable for long-term 3D imaging are desirable. Describe the included objective set (magnification, NA, working distance, immersion type). |
| Suppliers confirmation and description; |  |  |
| - 2 | B | A large field of view (FOV) is highly desirable to reduce or avoid stitching. Ideally, a single organoid (approximate diameter 250 µm) should fit within one FOV at 20x and 40x. Specify the FOV (µm and pixels) for relevant objectives and describe stitching options. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The system should allow fast imaging of full 96-well plates in confocal mode. Describe the time required to image a full 96-well plate with 1, 2 or 3 laser lines, using representative exposure times. |
| Suppliers confirmation and description; |  |  |
| - 2 | B | The confocal solution should provide well-resolved sub-cellular information in 3D samples (e.g. spinning disk or similar). Good optical sectioning in organoids and similar models is required. Describe expected resolution (X, Y, Z) and depth performance. |
| Suppliers confirmation and description; |  |  |
|  | B | Inclusion of additional imaging hardware solutions that improve image quality and/or resolution will be viewed very favourably. |
| Suppliers confirmation and description; |  |  |
| - B | B | The availability of advanced or smart acquisition (adaptive microscopy) strategies will be considered highly favourable. Please describe the available solutions, their ease of implementation, and the realistic training time required for users to learn how to use these options. |
| Suppliers confirmation and description; |  |  |
| - DETECTORS AND DATA QUALITY |  |  |
|  | A | The system shall be equipped with detectors suitable for both transmitted-light and fluorescence imaging. |
| Suppliers confirmation and description; |  |  |
|  | A* | The system shall be equipped with at least one high-sensitivity, fast detection system with a large FOV, suitable for high-throughput imaging. The detectors shall support image acquisition with a bit depth to 16 bits. |
| Suppliers confirmation and description; |  |  |
| - 3 | B | Included camera(s) should have high quantum efficiency (≈ 90 % or higher in 400–750 nm range), low read noise (≈ 1 e⁻ or lower) and high dynamic range. Describe the offered detector(s). |
| Suppliers confirmation and description; |  |  |
| - 3 | B | The camera pixel size should be well matched to the offered high-resolution objectives to allow proper Nyquist sampling. Specify effective pixel size (µm) at relevant magnifications. |
| Suppliers confirmation and description; |  |  |
| - 3 | A | The camera shall allow fast acquisition speeds (> 50 fps at full FOV). Describe achievable frame rates at relevant frame sizes. |
| Suppliers confirmation and description; |  |  |
| - 3 | B | The camera chip should have at least 2048 × 2048 pixels. State sensor size and FOV (in µm) with the offered objectives. |
| Suppliers confirmation and description; |  |  |
|  | B | The camera should use a high-speed data interface suitable for fast image acquisition and transfer. Describe the camera connection type(s) and their maximum supported data rates. |
| Suppliers confirmation and description; |  |  |
| - CONFOCAL MODULE AND LASER SOURCES |  |  |
| - 4 | A* | The system shall be equipped with a fast confocal module providing optical sectioning of samples. Describe the confocal technology (e.g. spinning disk) and its suitability for high-throughput 3D imaging. |
| Suppliers confirmation and description; |  |  |
| - B | B | Describe the process and ease for facility staff to exchange the disks in case several spinning disks are available. |
| Suppliers confirmation and description; |  |  |
| - 4 | A | The detection system should allow to image a large FOV. Specify the maximum FOV that can be scanned with representative objectives. |
| Suppliers confirmation and description; |  |  |
| - 4 | A | The camera should support fast imaging with a large field of view. Specify confocal imaging speeds for full and reduced frame sizes (e.g. 2048 × 2048, 1024 × 1024, 512 × 512). |
| Suppliers confirmation and description; |  |  |
| - 4 | A* | The confocal module should provide deep imaging (> 100 µm where relevant) with good rejection of out-of-focus light and high confocality. Describe achievable resolution (X, Y, Z) and section thickness with the offered objectives. |
| Suppliers confirmation and description; |  |  |
| - 4 | B | The confocal module should be compatible with fast sequential multichannel acquisition. Describe relevant features. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe features that minimize photo-damage to the sample (e.g. laser modulation, adaptive illumination). |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the solution for separation of laser excitation and fluorescence emission light (e.g. dichroics, emission filters). |
| Suppliers confirmation and description; |  |  |
| - 4 | B | All lasers should be solid-state or similar technologies, with long lifetime, low maintenance and sufficient power for fast imaging. Describe expected durability and minimum guaranteed power output per laser. |
| Suppliers confirmation and description; |  |  |
| - 4 | B | The system should be upgradable with additional lasers; if field-upgrades are not possible, provide options for alternative laser sources to cover additional fluorophores (e.g. CFP, YFP, infrared dyes). |
| Suppliers confirmation and description; |  |  |
| - OBJECTIVE CONSTELLATION OPTIONS |  |  |
|  | A* | The objective set must be suitable for imaging live and fixed cells, tissues, organoids (150–400 µm depth) and small organisms (e.g. Drosophila, zebrafish). The offer must include at least one high-quality high-magnification objective (e.g. 40x–60x) suitable for 3D imaging of organoids and similar samples. |
| Suppliers confirmation and description; |  |  |
|  | A | List, detail and price objective options that are well suited for the models described above. Options should include low-magnification (≈ 4x) and higher-magnification objectives (≈ 40x–100x). If possible, include high-end water or silicone-immersion objectives (minimum working distance 300 µm) or suitable immersion lenses; specify effective NA in different immersion media. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the maximum number of objectives that can be mounted in the turret simultaneously. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Detail DIC equipment for the offered objective lenses where relevant. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Objectives should have a high degree of chromatic aberration correction suitable for confocal imaging. Describe the wavelength range over which the objectives are corrected. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Objectives should be corrected (flat field and chromatic aberration) over a large FOV. Specify the corrected field size. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Objective constellation should include at least one objective with long working distance and one high-quality high numerical aperture objective; Specify for each objective. |
| Suppliers confirmation and description; |  |  |
| - 5 | B | Provide a list of additional dry, water, oil and silicone-immersion objectives suitable for the system and anticipated user needs. |
| Suppliers confirmation and description; |  |  |
| - B | B | Describe the ease of use and maintenance of different objectives, particularly in cases where user interaction is required (e.g. correction collar adjustment, use of a water dispenser). |
| Suppliers confirmation and description; |  |  |
| - SAMPLE MOUNTING AND CLIMATE CONTROL |  |  |
| - 6 | A | The system shall be equipped with a heated cage incubator (or similar) providing stable environmental control suitable for extended live-cell imaging experiments. |
| Suppliers confirmation and description; |  |  |
| - 6 | A | The system shall include the necessary equipment to control temperature, CO₂ and humidity. |
| Suppliers confirmation and description; |  |  |
| - 6 | A | Environmental control units shall provide accurate and active temperature control, digital/active CO₂/air mixing and active humidity control. Objective heaters should be available if deemed necessary. Describe the proposed solutions. |
| Suppliers confirmation and description; |  |  |
| - 6 | A | The CO₂ controller shall be an active mixing unit capable of mixing 100 % CO₂ with ambient air to the desired CO₂ concentration. A simple flow-control unit is not sufficient. |
| Suppliers confirmation and description; |  |  |
|  | B | Additional environmental control capabilities (e.g. hypoxia) will be considered favourably. |
| Suppliers confirmation and description; |  |  |
| - B | B | Describe the maintenance requirements of the environmental control system, if any. |
| Suppliers confirmation and description; |  |  |
| - 6 | B | The heated system should include sample holders compatible with several sample formats (e.g. multi-well plates, smaller multi-well imaging chambers, individual 35 mm dishes). The system shall include a slide holder accommodating conventional glass slides (76 mm × 26 mm). Flexibility to accommodate non-conventional sample format dimensions (like custom-made chips) will be viewed positively. |
| Suppliers confirmation and description; |  |  |
| - HIGH-THROUGHPUT AND PLATE HANDLING |  |  |
| - 7 | A | The system shall support high-throughput imaging of multi-well plates (e.g. 96-wells), with precise plate-based coordinate handling and plate-map definitions. Describe supported formats and mechanical constraints. |
| Suppliers confirmation and description; |  |  |
| - 7 | B | The stage or plate-handling solution should provide high accuracy and repeatability suitable for long-term point revisiting in all wells of multi-well plates. Describe positioning accuracy and repeatability specifications, including performance when the plate is removed from the instrument (e.g. for treatment) and subsequently reinserted. |
| Suppliers confirmation and description; |  |  |
| - 7 | B | Describe any high-content/high-throughput analysis workflows integrated in the system software that are particularly suited for plate-based experiments. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe compatible plate geometries for high-throughput (e.g. round or flat bottom, low or high profile plates). |
| Suppliers confirmation and description; |  |  |
| - B | B | The system should include flexible autofocus solutions suitable for high-throughput, plate-based imaging. Describe available autofocus methods, their impact on speed, photobleaching, and robustness, as well as their compatibility with different vessel types. |
| Suppliers confirmation and description; |  |  |
| - SOFTWARE, AUTOMATION AND ANALYSIS |  |  |
| - 8 | A* | The acquisition workstation shall come with dedicated software for multi-dimensional image acquisition and processing. The software should integrate all relevant hardware and support fast, hardware-synchronized imaging operations. |
| Suppliers confirmation and description; |  |  |
|  | A | Software controlling all parts of the system (microscope, stages, climate control, illumination, detectors, etc.) shall be included. |
| Suppliers confirmation and description; |  |  |
| - 8 | B | The software should allow automated multi-dimensional experiments (channels, z, time) at multiple positions and for different plate formats, both for widefield and confocal imaging. |
| Suppliers confirmation and description; |  |  |
| - 8 | A | The software should allow for fast setup of experiments. Describe how experiment setup is supported (e.g. templates, experiment editors, plate layout tools). |
| Suppliers confirmation and description; |  |  |
| - 8 | A* | The software shall include user-friendly setup dialogs (e.g. wizards) for multi-dimensional image acquisition, stitching and multi-position imaging. Please describe solutions, their ease of implementation, and the realistic training time required for users to learn how to use these options. |
| Suppliers confirmation and description; |  |  |
| - 8 | A | The software shall be capable of supporting high-content screening experiments. Describe this functionality and how such experiments are configured. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the available analysis modules; the level of analysis they provide (e.g. mask images, CSV tables, plots, reports); the types of image data they can handle (e.g. individual planes, Z-stacks, projections, time series, 3D renderings); and whether quantitative values can be directly visualized and contrasted with the corresponding images. |
| Suppliers confirmation and description; |  |  |
| - A | B | Describe the file format(s) and folder structure used by the system, and how they are adapted to high-throughput imaging workflows. |
| Suppliers confirmation and description; |  |  |
| - A | B | Describe how metadata are stored and handled (e.g. experimental and imaging settings, plate maps, analysis results) and how easily they can be searched, retrieved, and exported for downstream use. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the solution for plate visualization (e.g. synchronized lookup tables across wells, consistent min/max intensity settings, time-lapse data visualization, overview heatmaps, and other tools that facilitate comparison of images and signals across the plate). |
| Suppliers confirmation and description; |  |  |
| - 8 | B | The software should allow smart microscopy (i.e. altering imaging parameters based on feedback from automated image analysis during acquisition). Describe the available possibilities and user configuration options, as well as their ease of implementation and the realistic training time required for users to learn how to use these options. |
| Suppliers confirmation and description; |  |  |
| - 8 | B | The system should include customizable and easily implemented analysis pipelines for common assays relevant to the described models. Describe included analysis modules and options for user-level customization. |
| Suppliers confirmation and description; |  |  |
|  | B | The system should include analysis options that make use of AI or machine learning (e.g. for segmentation, classification, object detection, or anomaly detection). Describe these AI-based analysis functionalities and how they can be configured and used by end users. |
| Suppliers confirmation and description; |  |  |
|  | B | The inclusion of additional licences for analysis workstation(s) will be viewed highly favourably. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe any solutions for importing images from third-party instruments for analysis, if available. |
| Suppliers confirmation and description; |  |  |
| - OPTIONS AND UPGRADES |  |  |
| - 9 | B | To anticipate and accommodate developing user needs, modularity and future upgradability with well-integrated modules will be viewed highly favourable. Include a list of current and expected (2-year horizon) hardware and software upgrades, how they integrate with the offered system and the complexity of installation (duration, on-site vs factory, compatibility). |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Alternative and additional cameras suitable for the system, including options with smaller pixel size. Provide options and prices. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Differential interference contrast (DIC) solutions. Describe possible DIC configurations and associated components. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Fast z-scanning solutions such as piezo stages. Describe available options, integration and pricing. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Warranty extension. Give prices per year and indicate for how long the warranty can be extended. |
| Suppliers confirmation and description; |  |  |
| - 9 | B | Preventive maintenance contract (not covering parts). Give prices per year and indicate for how long a maintenance contract is available. |
| Suppliers confirmation and description; |  |  |
| - WORKSTATION AND SOFTWARE |  |  |
|  | A | The offer should include a high-end workstation suitable for the system, with appropriate memory, fast temporary storage and long-term storage for large time-lapse datasets and high-frame-rate imaging. Detail the minimal specifications required to guarantee full system functionality. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The offer should include one large or two medium-sized monitors that are ergonomic and have high resolution. Detail the minimal specifications required of such screens. |
| Suppliers confirmation and description; |  |  |
|  | B | The acquisition workstation should come with the software described in Section 8 and must be configured to support hardware-synchronized multi-dimensional imaging. |
| Suppliers confirmation and description; |  |  |
| - ANTI-VIBRATION TABLE |  |  |
| - 1 | B | Describe options and prices for active anti-vibrational table including silent compressor if one is required for the instrument. |
| Suppliers confirmation and description; |  |  |
|  | B | Future upgrades to the system (i.e. additional optical components or modules) should fit on the anti-vibration table. Please state available size options (dimensions, load capacity, configurability) and dimensions required to incorporate future large hardware extensions. |
| Suppliers confirmation and description; |  |  |
| - DELIVERY AND SERVICE |  |  |
| - 1 | A | Delivery should not exceed 6 months after received written order. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Indicate expected delivery time for the equipment as well as time requirements for installation until the equipment is ready for use. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | UiO has limited space for long-term storage and no specialized equipment for installation or internal transport. Specifically indicate any such needs. |
| Suppliers confirmation and description; |  |  |
|  | B | Provide realistic and maximum estimates for delivery time for spare parts, and specify. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Provide and specify realistic and maximum timelines for repair routines, from reception of a request from the UiO facility to on-site repair. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe the contact line for technical support (phone, e-mail, general portal, etc.) and expected response times. This service should be free of charge for hardware and software throughout their lifetime. |
| Suppliers confirmation and description; |  |  |
|  | A | Include a list with prices for parts required for preventive maintenance visits, as well as spare parts. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | The provider will guarantee access to user support, technical support, spare parts and full repair for at least 7 years. Indicated guarantees beyond 7 years will be evaluated positively. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Describe, in short, the maintenance routines (estimated duration and frequency) and equipment needed to be followed by the facility manager as part of local preventive maintenance routines. |
| Suppliers confirmation and description; |  |  |
|  | B | Describe nature, content, frequency and duration of the preventive maintenance visits scheduled during the warranty period. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | All software updates needed to run the system should always be available and free of charge throughout the lifetime of the software and system. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Describe the organization, size and availability of user and technical support. Indicate the location of the closest technical support facility. Presence of local expert support will be evaluated positively. |
| Suppliers confirmation and description; |  |  |
| - WARRANTY |  |  |
|  | A | Describe the warranty solutions included in the offer and alternatives if any. Indicate the guaranteed running time (in % of working days) during the warranty period. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | Warranty must comply with Norwegian regulations and cover costs for all repairs, hardware and software needed during this process. It should also cover all costs for travel and working fees /during these visits. |
| Suppliers confirmation and description; |  |  |
| - 1 | A | During the warranty period, preventive maintenance visits (PMVs) must be included and without added charges. The last PMV will be near the end of the warranty and include a software update. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | List any parts that are covered by a longer warranty. |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Provide an overview and pricing of service/support solutions after expiration of the warranty. |
| Suppliers confirmation and description; |  |  |
| - TRAINING |  |  |
|  | A | Our facility serves a large community with varying user expertise. At least 2 days of on-site user training for at least 4 users must be included. |
| Suppliers confirmation and description; |  |  |
|  | A | Additional training sessions for advanced users (up to 4 users) must be included. Describe the content and duration of such sessions. |
| Suppliers confirmation and description; |  |  |
| - ENVIRONMENTAL FOOTPRINT |  |  |
| - 1 | B | Indicate power consumption (in kWh) at normal use (10h/day). |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Indicate the extent to which equipment components can be recycled upon disposal (in %). |
| Suppliers confirmation and description; |  |  |
| 15.3 | B | Indicate the expected normal and maximum allowed noise level (in dB) of the entire system (including climate-control setup) during normal use. |
| Suppliers confirmation and description; |  |  |
| 15.4 | B | Solutions that incorporate/recycle existing surplus MIP equipment will be reviewed positively. These elements include objectives (Nikon MRD01691, MRD01905, MRH01401, MRH48430; Zeiss 420782-9900-799, 421787-9970-799, 421767-9970-000; Olympus 1-U2B836, 1-UB827, 1-U2C377), antivibration tables (120x85 or 200x150, both with metric threads), or microscope stands (Nikon Ti2-Eclipse). Describe compatibility, mode and extent of incorporation, and limitations. If needed, further details regarding equipment can be provided upon request. |
| Suppliers confirmation and description; |  |  |
| 15.5 | B | The provider will guarantee access to user support, technical support, spare parts and full repair for at least lifetime. Indicated guarantees beyond lifetime years will be evaluated positively. |
| Suppliers confirmation and description; |  |  |
| 15.6 | A* | State the guaranteed life-time period, at least 7 years, of the system which include available spare parts and software upgrades |
| Suppliers confirmation and description; |  |  |
| - ADDITIONAL CONSIDERATIONS |  |  |
|  | A | All equipment must meet the relevant directives and standards. Include separate documentation that supports this (Declaration of Conformity). |
| Suppliers confirmation and description; |  |  |
| - 1 | B | Describe the room requirements for installation and normal operation of the equipment (consider vibrations, power outlets, temperature control, running water and gas). |
| Suppliers confirmation and description; |  |  |

Appendix 3 Price and payment terms

# Price

[All prices and specific payment terms for the Supplier's deliverables shall be stipulated here. Any agreed deviations from the agreement’s principle stating that the amounts are in Norwegian kroner (NOK) and that prices are exclusive of Value Added Tax, but inclusive of customs duties and any other indirect taxes, shall be specified here]

- All the Buyer's payment obligations appear from this appendix, as well as Appendix 3a Price Matrix

Table 2: Payment plan

Invoice shall not be issued before delivery day.

# Price adjustment

- Fixed price period

Prices shall remain fixed during the first 12months from the date of contract signing.

- Access to price adjustment

Both parties may submit a request for price adjustment once per calendar year. The first price adjustment may take effect only after the expiry of the fixed price period. Subsequent price adjustments may take effect no earlier than 12 months after the previous price adjustment.

- Calculation of price adjustment

The price adjustment shall be based on the change in Statistics Norway’s Consumer Price CPI – Main Index over the most recent 12 available months, calculated from the time the price adjustment request is submitted.

- Efficiency and technological savings

If the supplier achieves cost savings as a result of efficiency improvements, process improvements, or technological developments, such savings shall be taken into account in the price adjustment.

Any expected efficiency gains, as specified in the contract appendices, shall be deducted from the calculated price change. The supplier shall, upon request, be able to document efficiency measures and the related financial benefit. The Contracting Authority has the right to require that such savings result in a price reduction.

## Invoice format

All UiO’s European suppliers must send invoices and credit notes in Peppol-format. UiO’s Peppol ID: 0192:971035854.

Foreign suppliers who are not able to send invoice in Peppol-format, may send PDF invoices to invoices@admin.uio.no

The Supplier must use UiO's chosen format and invoice processing method at all times. At least three (3) months' notice must be given before any change takes place. The Supplier carries their own costs associated with the delivery of electronic invoice as well as credit notes.

## Invoice quality

Invoices must have a reference to the correct purchase order number from the Buyer’s purchasing system. The Buyer reserves the right to return all invoices without correct reference.

The Buyer uses a system for analysis and statistics, which demands high quality of the inbound invoices. The invoices sent to UiO must therefore contain following information on each line row:

- Unique product number pr. product name

- Full product description

- Serial number/ unique item number (applies to purchase of equipment with serial number)

- Product price (excl. VAT)

- VAT rate

- Alternatively, any agreed fees (example: minimum order fee)

- VAT registration number and/or tax ID number placed at the invoice header

The product number and product name on the invoice must correspond to what is stated in the catalogue /punch out. To achieve auto match of invoices, it is necessary that the product number in both the catalogue /punch out are identical.

To ensure that invoices are processed correctly via UiO's systems, a test invoice must be conducted. Please contact the officer responsible for the contract and make further agreements.

Person sensitive information, such as health information, ethnicity or other sensitive personal data covered by GDPR Article 9, shall not appear in the invoice lines.

Please follow the link to guide on implementation of electronic invoice at UiO: https://www.uio.no/english/about/collaboration/invoice/index.html

Appendix 4 Administrative provisions

Project- and progress plan

[Fill out the relevant deliveries, partial deliveries and activities with the associated schedule, commencement date and delivery date for the deliverables]

Key personnel

The following Supplier’s key personnel shall carry out the delivery of the service/goods:

| Name | Position | Area of expertise | E-mail | Telephone |
| --- | --- | --- | --- | --- |

Approved subcontractors

The Buyer approves the following subcontractors:

[To be completed at the contract formation]

Use of third parties

The Buyer will involve the following third parties:

[To be completed at the contract formation]

Delivery times

Requirement stated in Appendix 1 and answered in appendix 2

Place(s) of delivery

Sognsvannsveien 9, Domus Medica,0372 Oslo

Routines for order acquisition, handling and follow-up

The Supplier must be able to receive orders from the Buyer’s purchasing system. Order confirmation must be sent to the stated order reference (on the order) and must contain the following information:

- Name of the recipient

- Order description (product and quantity)

- Order number (UiO's reference number starting with number 6)

- Delivery time

- Any partial deliveries

Delays or other deviations associated with the delivery must be reported immediately to the Buyer, once the deviation is discovered.

Information on contract, contractor and employees

It is the client’s duty to report the assignment when the value of the assignment exceeds NOK 20.000. The foreign contractor is obliged to report the employees that work on the assignment.

Further information on the reporting obligation can be found on the Norwegian Tax Agency's website:

https://www.skatteetaten.no/en/forms/rf-1199-information-about-contracts-contractors-and-employees/

VAT and Registration Requirements for Foreign Suppliers

The Supplier acknowledges that Norwegian VAT rules differ from general EU VAT rules. In Norway, foreign suppliers delivering services physically in Norway (such as installation, training, or on-site support) — either alone or in combination with goods — and where the total value of such deliveries to one or more Norwegian customers equals or exceeds NOK 50 000 within a 12-month period, are required to register in Norway. The reverse charge mechanism commonly used within the EU does not apply in these cases.

The Supplier shall, prior to issuing any invoice, be registered in Norway with a Norwegian organization number (NUF) and be registered in the Norwegian VAT Register. The Supplier shall ensure that all invoices issued under this Contract include and handle Norwegian VAT in accordance with applicable Norwegian tax regulations.

The Customer shall have the right to withhold payment of any invoice until the Supplier has fulfilled the above registration requirements. This applies regardless of whether the Contract is entered into with the foreign parent company or a Norwegian-registered branch.

For more information and guidance on what foreign suppliers must do, see:

- https://info.altinn.no/en/start-and-run-business/direct-and-indirect-taxes/indirect-taxes/vat-specific-information-for-foreign-companies/

- https://info.altinn.no/en/start-and-run-business/planning-starting/Choosing-Legal-Structure/norwegian-branch-of-a-foreign-company-nuf/

- https://www.brreg.no/en/other-types-of-organisation/norwegian-registered-foreign-business-nuf/duty-to-register-with-the-register-of-business-enterprises-for-nuf/?nocache=1738149867038

- https://www.skatteetaten.no/en/business-and-organisation/foreign/registration-obligation/registration-obligations-for-foreign-limited-businesses-in-norway/

Appendix 5 Options

[All options relevant under this agreement shall be described in this Appendix (does not apply to option for contract prolongation). Contractual clauses concerning changes/options needs to state clearly which changes the Buyer can make, to which extent and under which terms. Such clause does not allow the contract’s overall nature to change].

Appendix 6 Changes to the general contractual wording

| Clause | Original text | New text |
| --- | --- | --- |

Appendix 7 Changes subsequent to the formation of the contract

The table below shall enlist all changes or additions to the contract subsequent to the contract formation.

The change table shall be kept by date on a continuous basis and state who initiated the change.

In the column “Date” the date of enforcement of the change shall be stated.

In the column “Initiated by” it shall be stated which of the parties requested the change.

In the column “Reference” the reference to the place in the contract where the change or amendment belongs shall be stated.

In the column “Change” the text changing the existing text shall be inserted.

In the column “Addition” the text that is added to the contract shall be stated.

In the column “Change” the following formatting shall be used:

Striked out text = deleted text

Bolded text = new text

| Date | Initiated by | Reference | Change | Addition |
| --- | --- | --- | --- | --- |