## 1. Guidance for use
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Project 1105601 UiO Livsvitenskap brukerutstyr
K935.10 Konfokal mikroskop
Tenderer:
HOW TO FILL OUT THE DOCUMENT
Please note that there are several tabs to be filled out.
Columns and rows are not to be deleted.
The tenderer shall only fill out fields with a white background. Fields of other colors are not to be used or altered by the Tenderer.
The tenderer shall answer YES or NO to the question on whether or not a requirement is fulfilled. All rows must be answered.
The tenderer shall answer the defined requirements. The answer shall have sufficient information for the contractor to carry out a full assessment of the tender without the need to open attachments. If the tenderer needs to provide additional or more detailed information than this document allows for, it is allowed to refer to attachments. In case of reference to attachments, the tenderer shall provide the corresponding attachment nr and relevant pages.
Minimum requirements (A-criteria) are not related to an award criterion. Lack of fulfillment of minimum requirement can lead to rejecton of the tender, cf Tilbudsinvitasjon (Tender invitation).
B-criteria are related to an award criterion. The tenderers answer to B-criteria are the basis of the contractors evaluation and assessment. B-criteria are defined as B1, B2 or B3 criteria in the corresponding column for each criterion. The definitions indicate the assigned weight of the award criterion in the tender evaluation. Award criteria indicated as B3 are assigned most weight, while the B1 criteria are assigned the least weight.
Information criteria (I-criteria) are solely for informative purposes, and are not part of the evaluation.
| Valuta: | Lisens:
| NOK | Systemlisens
| SEK | Brukerlisens
| DKK | Utstyrslisens
| EUR | Analyselisens
| USD | Annet
GBP

## 2. Point Scanning Confocal
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Navn på artikkel | Art. 025.02.001:3 Konfokalmikroskop, Point Scanning Confocal, opprett
Criterion nr | Description of criterion | Criterion type | Supplementary information/Additional information | Fulfillment of criterion Yes/N0 | Supplier response
1 | QUALITY AND FUNCTIONALITY CRITERIA
1.1 | The equipment shall be CE marked in accordance with current directives and fulfill all relevant Norwegian laws, standards, requirements and norms for the equipment offered. Declaration of conformity for CE marking (Declaration of conformity) shall be attached to the offer. | A
1.2 | Intended use: The system shall be a state-of-the-art multimodal, fully automated, upright, point scanning confocal microscope with high optical resolution, sensitivity, and imaging speed. It shall enable detailed visualization and quantification of subcellular structures, as well as imaging of small animals (e.g., zebrafish, C. elegans) and tissues. Fast scanning is required for non-invasive live 3D/4D imaging, supported by accurate data acquisition and analysis across a wide range of biomedical applications. | A | The microscope shall provide modular flexibility, intuitive software integration, and future upgrade options to accommodate evolving research needs. Intended samples: Live cells, organoids and embryos, thick tissue sections, live tissue slices, fixed specimens, cleared samples, as well as opaque samples in material research.
1.3 | State the estimated delivery time from when the order is received from the customer until the equipment is delivered. | A
1.4 | A complete user manual and safety data sheet in a Scandinavian language or English shall accompany the delivery, preferably in electronic format. | A
1.5 | The Tenderer shall, in addition, submit a detailed breakdown of all components included in the main offer, presented in separate line items including pricing, as a separate attachment. This is intended to provide an overview of all components included in the main offer. It does not replace Appendix 3.0 – Price Form. | A
Microscope Frame and Sample Environment
1.6 | The instrument shall be a fully motorized upright point-scanning confocal microscope with a motorized Z-drive and nosepiece. | A
1.7 | The microscope shall be equipped with an Epi-fluorescence module so the sample can be examined via the oculars and camera. | A
1.8 | The offer shall include a dedicated high-performance anti-vibration table for the confocal microscope, as well as an additional table to support the ancillary equipment required for its operation. | A
1.9 | The system shall include a fully motorized XY stage for high-speed scanning. | A
1.10 | Please specify the travel range, speed, and precision in x, y and z of the stage. | I
1.11 | The microscope shall be equipped with several sample holders specifically for 35 mm imaging dishes, and slides. | A
1.12 | The system shall have an autofocus. | A
1.13 | The supplier should offer a cage box incubation and environmental control system, including temperature and humidity regulation, as well as a CO₂ mixer. This critria is evaluated on whether the Supplier can offer an incubation system or not. The option shall be priced separately in Appendix 3.0 – Price Form K935.10, under tab 4 – Tilvalg. The option price will not be subject to evaluation | B1
1.14 | The offered solution should have additional gas control possibilites, including but not limited to oxygen (O₂) control for hypoxia simulation. Please decribe the possible solution. This criterion will be evaluated based on whether such additional gas control capabilities are available. If such capabilities are available, they shall be priced separately and listed as individual line items in Appendix 3.0 – Price Form K935.10, under tab 4 – Tilvalg. | B1
General Optics and Visual Inspection Modalities
1.15 | The instrument shall include 6 chromatically corrected APO objectives: 10x Air, 20x Air, 20x Water Dipping, 40x Water Dipping, 60x/63x Oil and 60x/63x Silicon/Glyserol. | A
1.16 | Ref. point 1.15: High Numerical Aperture (NA) and long Working Distance (WD) are considered advantageous. The criterion will be evaluated based on the user’s assessment of the documented objective functionality and performance, with particular emphasis on NA and WD. | B3
1.17 | The tenderer shall provide a list of objectives not included in item 1.15, that are considered suitable for the intended functions/use. These objectives may differ forme those listed in point 1.15 in term of magnification, immersion type, workring distance, NA-values, etc. The objectives shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 4. Tilvalg. The price will not be subject to evaluation. | A
1.18 | The microscope shall be equipped with differential interference contrast (DIC), phase contrast or other equilvalent bright field contrast method. | A
1.19 | The system should be capable of imaging opaque samples (reflected light imaging). The criterion will be evaluated on a binary basis (yes/no), depending on whether the system is capable of imaging opaque samples. | B1
1.20 | The microscope shall be equipped with a monochromatic camera optimized for epifluorescence imaging, with a minimum Quantum Efficiency (QE) of 80%, and a color camera for epifluorescence and bright field imaging. | A
1.21 | The microscope should have a large optical field of view (FOV) at the camera port. Please state the optical field of view at the camera port. The microscope with the highest FOV at the camera port will get the highest score, and microscopes with smaller FOV get a lower score, calculated linearly from the maximum. | B2
1.22 | The system should have a large effective field of view. Please specify the maximum achievable Field of View (FOV) using the offered 10x dry objective at an excitation wavelength of 488nm. Evaluation will be based on the diameter (mm) of the usable imaging area. The system with the largest documented usable imaging area will be awarded the highest score. | B3
Excitation and Scanning
1.23 | The system shall include 4 lasers centered around the wavelengths 405nm, 488nm, 561nm, and 647nm, or a single laser that covers the same wavelenghts. | A
1.24 | One additional laser line in the infrared scanning area (e.g. 685nm, 730nm, 750nm or 785nm) shall be offered as an option. The tenderer shall state the wavelenght of the offered infrared laser line. The additional laser shall be priced in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Utstyr Opsjon. All components required for full functionality shall be clearly specified and included in the option price. | A
1.25 | The instrument shall have a motorized laser point skanner supporting zoom, bidirectional and ROI scanning. | A
1.26 | The system shall an acheive acquisition speeds of at least 10 fps at a resolution of 1024 x 1024 pixels. | A
1.27 | Ref. point 1.26: The system will be evaluated based on acquisition speed (frames per second, fps) at a resolution of 1024 × 1024 pixels. An acquisition speed above 10 fps will receive a higher scores. The system with the highest acquisition speed will recieve the highest score. | B3
1.28 | The system shall be able to combine brightfield and confocal simultaneously, and brightfield with super-resolution simultaneously. Describe the offered solutions and specify the imaging speed in each of the simultaneous modes. | A
Detection and Resolution
1.29 | The system shall have at least 4 high-QE confocal detectors and one brightfield detector. | A
1.30 | Ref. point 1.29. Please specify the QE level and the spectral sensitivity of the detectors. Please specify the quantum efficiency (QE) level and spectral sensitivity of the detectors. Detector quantum efficiency curves and spectral sensitivity curves for the offered solution shall be attached. Higher quantum efficiency and broader spectral sensitivity will be scored higher. | B3
1.31 | The detectors should have a low level of dark noise. Please specify detector dark noise level. Solutions offering detectors with lower dark noise levels will be scored higher. | B1
1.32 | The transmitted light detector (T-PMT/BF detector) shall be capable of being positioned, or the system shall utilize a reflected-light detector, to enable visual inspection and digital capture of opaque samples. | A
1.33 | Please describe the efficiency of the light path (e.g., filter-based or prism-based dispersion). Please emphasize the system's ability to minimize spectral crosstalk through hardware design or software-based spectral unmixing during multi-channel acquisition. | I
1.34 | Detectors shall allow for spectral detection. | A
1.35 | The system shall be able to perform sub-diffraction (super resolution) imaging. The resolution must be better than 250nm. | A
1.36 | Ref. point 1.35. The system should provide high spatial resolution for super-resolution imaging. Please describe the underlying mechanism/technology for the offered super-resolution. The evaluation will be based on the users’ assessment of the proposed solution. The highest score will be given to the system with the lowest value in nm (x,y resolution). | B2
Advanced Modalities
1.37 | Multiple scanning modes, frame, line, and ROI scan is required. | A
1.38 | The software shall provide full support for FRET, FRAP, and automated tile-stitching. The software must seamlessly stitch the scanned tiles. | A
1.39 | The system shall be equipped with automated navigation tools (such as sample finder or sample overview), including the ability to utilize pre-defined sample templates for standardized multiwell plates. | A
1.40 | The offered solution shall include a comprehensive software package for advanced data aquisition and analysis. The sotfware shall include AI/deep learning functionality for image acquisition, image processing, and image analysis. The Tenderer shall provide a detailed description of the included functionalities. As a minimum, the drscription shall cover, but be limited to the follwing: - Data preprocessing - AI-supported analysis (live and post-acquisition) - Deconvolution/denoising - Segmentation, and data analysis The offered solution shall be delivered with complete user manual(s) for the software package. If the Tenderer refers to the user manual(s) to describe functionalities, in whole or in part, clear and specific references to the relevant pages covering the described functionalities shall be provided. | A
1.41 | This criterion looks at the offered software package as described by the Supplier in Point 1.40 and the functionalities that are included. The evaluation will be based on the users’ overall assessment of the functionality of the proposed solution/package, including its usability, performance, and relevance for the intended application. | B1
1.42 | Please describe the systems AI-driven or "smart" acquisition function and if the software can automatically identify regions of interest (ROIs) for high-resolution confocal scanning based on initial widefield detection. This requirement is for informational purposes only and will not be subject to evaluation. | I
1.43 | Please specify if there are any Optical or software correction for deeper sample scanning included in the solution. This requirement is for informational purposes only and will not be subject to evaluation. | I
1.44 | Please specify whether the instrument is open for future system upgrades (e.g. adding additional laser lines, new technology). Please also describe the upgrade capability of the system and list potential system upgrade and expansion in the options. This requirement is for informational purposes only and will not be subject to evaluation. | I
Two-photon options
1.45 | The system shall include an option for a two-photon laser system. The option shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Utstyr Opsjon. The stated price shall include all components, functionality and deliverables specified in points 1.45- 1.49. | A
1.46 | Please describe the offered two-photon laser system as specified in point 1.45, including its technical configuration, functionality, output power, and performance characteristics. Excitation of red fluorophores (approximately 1040–1064 nm) is preferred. The system will be evaluated based on the users’ assessment of the overall functionality of the pruposed solution, including tissue penetration, laser power, imaging depth, spatial resolution and software integration. Solutions demonstrating higher performance, greater capability, and better integration across these aspects will be awarded a higher score. | B3
1.47 | The offered solution shall include one Non-Descanned Detector 2 (NDD 2) -photon dectector. | A
1.48 | A NDD-2 detector with a high QE value is preffered. Please describe the offered NDD 2-photon detector and its QE value, ref. point 1.47. This criterion will be evaluated based on the detectors Quantum Efficiency (QE), where higher QE values will be awarded higher scores compared to lower QE values. | B2
1.49 | The offered option shall include one 20x/25x water immersion objective that is optimized for deep two-photon imaging. | A
1.50 | The Tenderer shall provide a list of optional objectives not included in point 1.49, that the Tenderer considers suitable for two-photon imaging. These objectives may have different magnification, working distance, NA values etc, than the objective described in 1.49. The listed objectives shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 4. Tilvalg. Prices will not be subject to evaluation. | A
1.51 | The system should utilize an optical design that minimizes the optical path length between the sample and the NDD detector. Please specify the optical design and the physical placement of the NDDs relative to the objective lens, including the distance (in millimeters or centimeters) from the back aperture of the objective to the NDD sensor. The detector configuration will be evaluated based on the physical optical path length between the sample and the detectors. Solutions with shorter optical path lengths will be awarded higher scores compared to solutions with longer optical paths lengths. | B3
1.52 | The two photon option should include the possibility for Fluorescence Lifetime Imaging Microscopy (FLIM). If FLIM is supported, please describe the underlying technology (time-domain or frequency-domain) and specify the acheivable lifetime measurement range. The proposed solution will be evaluated based on whether FLIM functionality is possible with the two-photon option. Please provide the price for FLIM solution in Appendix 3.0 – Prisskjema K935.10 under tab 4 - Tilvalg. The stated price must include all necessary hardware, detectors, electronics, and software required for a fully functional workflow. | B1
General Requirements
1.53 | The instrument and its associated transport packaging must fit through a door with dimensions of 90 × 205 cm (W × H). The tenderer shall describe the number of transport packages as well as their respective dimensions in cm and weight in kg. | A
1.54 | Please describe how the Supplier handles product updates or new software releases introduced between the time of tender submission and the actual delivery of the instrument. Please state whether upgrades to newer versions are included in the offered price, or if they entail additional costs. This requirement is for informational purposes only and will not be subject to evaluation. | I
IT Hardware
1.55 | The offered instrument shall include a dedicated workstation with a high-performance graphics card, high performance storage drives, a keyboard and a mouse. The instrument computer shall be delivered with all software modules necessary for instrument control. | A
1.56 | Two high-resolution monitors (27" or larger), or a single ultra-wide monitor providing an equivalent total screen area and display resolution shall be included in the offered solution. The Tenderer shall provide a technical description of the monitor(s) included with the instrument, explicitly specifying the screen size and display resolution. | A
References
1.57 | At least two relevant references for the offered solution shall be listed. The references shall relate to comparable installations. Each reference shall include: - Name of the customer/institution - Contact person - Telephone number and e-mail address - Description of the delivered solution and its scope The contracting authority reserves the right to contact the provided references. | A
Demonstration
1.58 | The supplier shall offer a demonstration of the offered solution, to be conducted during the evaluation phase of the tender. Please refer to Appendix 1.0 Invitasjon til anbudskonkurranse K935.10, point 4.4.1 for additional information on demonstrations, including the relevant time table. The demonstration shall be conducted as a physical demonstration at a suitable location. | A
1.59 | The demonstration of the offered solution, should be with the instrument set-up as described in this tender. The buyer have the right to use their own samples for the demonstration, should they wish it. The evaluation will be based on the users’ assessment of the proposed solution as shown in the demonstration, including user interface. Suppliers should showcase the following capabilities: - Live Sample Imaging: Water-dipping imaging of a live sample using at least two fluorophores, combined with brightfield and multi-point (XYZT) acquisition. - 3D Deep-Tissue Scanning: A 3D Z-stack of a thick sample to demonstrate the effective working distance and performance of the water-dipping objectives. - Large-Area Tile Scanning: A multi-channel tile scan of a large sample (minimum 20 fields of view) to demonstrate stitching quality and illumination uniformity. - Super-Resolution Mode: A demonstration of the system's super-resolution capabilities in terms of spatial detail and acquisition speed. - Image Processing Software: A demonstration of the software suite, specifically focusing on data handling, visualization, and analysis. -Optional: 2-Photon and FLIM: Performance demonstration of the 2-Photon excitation depth and the Fluorescence Lifetime Imaging (FLIM) module is preffered but not required. If a physical demonstration of any of these modes on the actual systems is not possible, a powerpoint (or similar) presentation containing technical details of these modes must be included instead. | B3
Training requirements
1.60 | Necessary super user training shall be offered for approx. 5 users. The training must be carried out by agreement. The traning is done on-site at the customer site. The training must ensure proper, safe and correct use of the equipment. The Vendor must anticipate that training will take place after takeover. Describe scope and content of the training. Training must be included in the tender at no additional costs. | A
1.61 | Please describe the capability and availability for providing on-demand training sessions for new or advanced users after the initial system installation period. This requirement is for informational purposes only and will not be subject to evaluation. | I
ICT requirement
1.62 | All mandatory program and security updates for both software and equipment shall be delivered and installed free of charge until the end of service. By update is meant improvement and "bugfix" of software. But not new features. | A
2 | SERVICE. OPERATION AND MAINTENANCE CRITERIA
Service requirement
2.1 | The Tenderer shall list prices for the following in Appendix 3.0 –Prisskjema K935.10 under tab 3. servicetjenester: - Attendance price, the price shall include all travel expenses and time spent per travel to the service point - Hourly price for the performance of service, covers time spent with the client or supplier The rates/prices will be included in the evaluation of the total cost of the tender. | A
Maintenance agreement
2.2 | An option for a maintenance agreement Level 1 Preventive shall be offered. The maintenance agreement shall include: - Preventive maintenance according to the manufacturer's specifications, including all wear parts/service kits/components/batteries that the manufacturer recommends replacing in connection with maintenance, including parts, all travel and accommodation expenses and any application follow-up or similar. - Documentation of work carried out (service report) shall be delivered to the client within 5 working days - Phone support The supplier shall schedule in advance the time for the maintanance with the Buyer. The agreement shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Vedlikeholdsavtaler opsjon | A
2.3 | An option for a maintenance agreement Level 2 Comprehensive shall be offered. - Comprehensive agreement with annual costs for preventive maintenance including parts, emergency calls, including parts, all travel and accommodation expenses and any application follow-up or similar. The agreement shall cover all breakdown repairs and parts, including lasers. - Documentation of work carried out (service report) shall be delivered to the client within 5 working days - Phone support The supplier shall schedule in advance the time for the maintanance with the Buyer. The agreement shall be priced separately in Appendix 3.0 – Prisskjema K935.10 under tab 3 - Vedlikeholdsavtaler opsjon. | A
Service and maintenance
2.4 | Please state the maximum delivery time from order to delivery at the customer's place, including delivery from "external warehouse". (1 calendar day = 24 hours). This requirement is for informational purposes only and will not be subject to evaluation. | I
2.5 | A complete service manual in electronic format shall accompany the delivery, if available. The service manual shall be in English or a Scandinavian language. | A
2.6 | Phone support shall be available during normal working hours. Communication with customer support shall be carried out in a Scandinavian language or English. Ordinary working hours are defined as weekdays at 08:00-16:00. Describe the support available and any conditions or limitations to this service. | A
2.7 | Please give a brief description of the procedure for carrying out preventive maintenance according to the manufacturer's instructions. Please attach protocols for carrying out preventive maintenance, as well as information on frequency and expected time consumption. This requirement is for informational purposes only and will not be subject to evaluation | I
2.8 | The offered warranty period for the system shall be stated. | A
2.9 | An option for an Extended Warranty Agreement shall be offered. The agreement shall be priced in 3.0 Prisskjema K935.10 under tab 3 - Garanti opsjon. | A
2.10 | During the warranty period, the Tenderer shall provide preventive maintenance at Service Level 1 (refer to point 2.2) at no additional cost to the Buyer. Periodic maintenance shall be carried out in accordance with the manufacturer’s recommendations. The final scheduled maintenance shall be performed within the last two months of the warranty period.. | A
2.11 | At the end of the warranty period, the delivered equipment shall have the latest available software version installed. No additional costs shall apply for this. | A
2.12 | If equipment-specific special tools, calibration devices, software, passwords, service keys, or similar are required to perform technical troubleshooting, maintenance, and/or emergency repairs, this shall be clearly described and priced, this shall be clearly described and priced in Appendix 3.0 – Prisskjema K935.10 under tab 4. Tilvalg | A
2.13 | The Tenderer shall describe which parts are expected to be replaced during the equipment’s expected lifetime, including the estimated frequency of replacement. Examples may include, but are not limited to, light sources, moving parts, and other consumable or wear components. | A
2.14 | Accsess to spare parts shall bed guaranteed for at least 7 years after delivery of the equipment. | A
2.15 | Access to spare parts beyond 7 years is emphasised. Please state the total number of years for which access to spare parts is guaranteed. | B3

## 3. Technical specifications
TECHNICAL SPECIFICATIONS
Statsbygg art. nr | Article name | Tenderers offer/model | Acoustic noise: (dB (A)) | Heat emission: Standby / operation (kW) | Room requirements: Temperature (oC incl peak-to-peak limits) Relative humidity min/max (% rH) | Power supply: (power plug, fixed plug, length of power cord) Effect (kW) Voltage (V) Ampere (A) Require dedicated circuit Require dedicated earthing | Specific electricity requirements: Emergency power supply UPS | Ethernet: Socket position, cable type, cable routing, plug type, other, capacity Any max cable length | Lighting source Cable routing | Other:
025.02.001:3 | Konfokalmikroskop, Point Scanning Confocal
Statsbygg art. nr | Article name | Tenderers offer/model | Couplings to: Gases (incl compressed air) Pressure (bar) Purity (%) | Other:
025.02.001:3 | Konfokalmikroskop, Point Scanning Confocal