## 1. Guidance for completion
REQUIREMENTS SPECIFICATION
Prosjekt 1105601 UiO Livsvitenskap brukerutstyr
K936.04.01 Automated pipetting and sorting system for zebrafish embryos, eggs and larvae
Supplier:
Guidance for completion
Please note that there may be multiple tabs to complete.
Columns and rows in the sheet must not be deleted.
The tenderer shall only write in fields with a white background. Fields that are colored shall not be completed or changed.
The tenderer shall answer YES or NO as to whether the requirement is met. All lines must be answered.
The tenderer shall respond to the requirements. The response shall contain sufficient information to enable the contracting authority to carry out a complete evaluation of the tender without having to open attachments. If the tenderer needs to provide more detailed information than there is room for in the requirements specification, reference may be made to an attachment. When referring to an attachment, the attachment number and page number must be specified.
Minimum requirements (A-requirements) are not related to an award criterion. Failure to meet minimum requirements may result in rejection, cf. the Invitation to Tender.
Evaluation requirements (B-requirements) are related to an award criterion. Responses to evaluation requirements will form the basis for the contracting authority's evaluation and requirement-setting. Evaluation requirements are abbreviated to B1, B2, or B3 in the column for requirement type. The abbreviation indicates the weight the evaluation requirement will have in the evaluation of the tenders. Evaluation requirements marked B3 have the highest weight, while those marked B1 have the lowest weight.
Information requirements (I-requirements) are for information only and will not be included in the evaluation.
| Valuta: | Lisens:
| NOK | Systemlisens
| SEK | Brukerlisens
| DKK | Utstyrslisens
| EUR | Analyselisens
| USD | Annet
GBP

## 2. Requirements Specification.
REQUIREMENTS SPECIFICATION
Requirement number | Description of requirement | Requirement type | Additional information | Requirement fulfillment Yes/No | Supplier’s response
QUALITY AND FUNCTIONALITY
1.1 | The equipment shall be CE‑marked in accordance with applicable directives and shall comply with all relevant Norwegian laws, standards, requirements and norms for the offered equipment. A Declaration of Conformity for the CE‑marking shall be attached to the tender. | A
1.2 | The instument must at a minimum be suitable for its intended use: The instrument is intended for use in research and laboratory environments for the automated handling, pipetting, transfer and sorting of live embryos, eggs and larvae, primarily of zebrafish (Danio rerio). The instrument is intended to reduce manual handling, improve consistency, and support standardized research protocols while minimizing mechanical stress and potential harm to living specimens. | A
1.3 | The estimated delivery time from receipt of purchase order to delivery of the equipment shall be stated. | A
1.4 | Complete user documentation in English shall accompany the delivery, preferably in electronic format. The customer shall, until the end‑of‑service date as defined by the manufacturer, receive electronic updates of the user documentation unsolicited and at no additional cost, regardless of file format. | A
Requirements for the Equipment
1.5 | The instrument shall be capable of automated handling, pipetting and sorting of live zebrafish (Danio rerio) eggs, embryos and larvae, without the need for manual intervention during normal operation. Documentation demonstrating compliance with the requirement shall be provided | A
1.7 | The instrument shall be compatible with relevant standard laboratory formats, such as Petri dishes, microtiter plates, or equivalent containers commonly used in zebrafish research. | A
1.8 | The instrument shall be constructed in a manner that minimizes the risk of cross-contamination between samples through appropriate design measures. The instrument shall also support cleaning and maintenance in accordance with documented procedures provided by the manufacturer. The supplier shall provide documentation describing: - How the construction of the instrument contributes to reducing the risk of cross-contamination (e.g. fluid paths, pipette tips, contact surfaces). - The recommended cleaning and maintenance procedures, including frequency and required tools or consumables. | A
1.9 | The instrument should demonstrate a documented capability to ensure gentle and controlled handling of live zebrafish (Danio rerio) eggs, embryos, and larvae, with the objective of minimizing mechanical stress, shear forces, and the risk of injury or mortality during automated pipetting, transfer, and handling operations. The requirement applies in particular to handling of organisms within relevant size ranges and geometries, including elongated and delicate specimens. Please provide the following information: - A technical description of the principles, features, and mechanisms that contribute to gentle handling (e.g. flow control, aspiration and dispensing principles, tip design, pressure regulation) - Documented data, test results, or published references demonstrating post-handling organism viability and injury/mortality rates (e.g. % viability, % injury), for zebrafish (Danio rerio) or comparable aquatic organisms The evaluation will be based on the contracting authority’s overall assessment of the documented performance and supporting evidence related to gentle handling of live organisms. Particular emphasis will be placed on, but not limited to, the following aspects - Documented post-handling viability rates - Documented injury and mortality rates - Relevance of the test conditions to the intended use - Quality, relevance, and credibility of supporting data Solutions that offer gentle handling with a low risk of injury and mortality, high post-handling viability, and well-documented performance supported by relevant and reproducible data will receive a higher score. | B3
1.10 | The instrument should be capable of dispensing defined liquid volumes with sufficient accuracy and repeatability to enable accurate and repeatable volumetric dispensing for transfer of single zebrafish (Danio rerio) eggs, embryos, and larvae. The system should ensure consistent volumetric performance across repeated dispensing cycles, supporting precise positioning and transfer of individual organisms without unintended co-transfer, retention, or failed dispensing. Please provide the following information: - Specifications for dispensing accuracy and repeatability (e.g. stated accuracy, coefficient of variation (CV), and relevant test conditions, including volume range) - Information on minimum, maximum, and typical dispensing volumes relevant for single egg, embryo, or larva transfer - Description of technical solutions that ensure volumetric precision and repeatability (e.g. liquid handling principle, calibration methods, sensors, and control mechanisms) The evaluation will be based on the contracting authority’s overall assessment of the documented volumetric performance and supporting technical information related to dispensing accuracy and repeatability for single-organism transfer. Particular emphasis will be placed on, but not limited to, the following aspects:: - Documented dispensing accuracy at relevant volume ranges - Documented repeatability (e.g. coefficient of variation under relevant conditions) - Suitability of the volume range for single egg, embryo, and larvae transfer - Assessed risk of unintended co-transfer, retention, or failed dispensing Solutions demonstrating high dispensing accuracy and repeatability at relevant volumes, a suitable volume range for single-organism transfer, and low risk of co-transfer or failed dispensing will receive a higher score. | B2
1.11 | The instrument should support handling of a broad range of zebrafish (Danio rerio) developmental stages, from early embryos to later larval stages, without the need for extensive reconfiguration. The offered solution should enable transitions between developmental stages primarily through software settings or minor adjustments, rather than extensive mechanical changes or hardware replacement. Please provide the following information: - A description of which developmental stages are supported (e.g. embryos, early larvae, later larvae) - A description of the reconfiguration required when switching between developmental stages (e.g. software adjustments, consumables, hardware changes) - Information on the complexity, time, and operational effort required to perform such transitions The evaluation will be based on the contracting authority’s overall assessment of the documented flexibility and usability of the solution, Particular emphasis will be placed on, but not limited to, the following aspects: - Range of supported developmental stages - Degree of required reconfiguration (software settings versus consumables or hardware changes) - Complexity and operational effort when switching between developmental stages Solutions that offer broad support for developmental stages and high flexibility with minimal need for reconfiguration will receive a higher score. | B3
1.12 | The instrument should demonstrate high precision and reliability in the handling and transfer of single eggs, embryos and larvae. Please provide the following information: - Documented error rates related to handling and transfer (e.g. loss, damage, failed transfers, or unintended co-transfer). - Description of any built-in control, monitoring, or verification mechanisms that support precise and reliable handling (e.g. sensors, visual verification, detection of successful aspiration and dispensing). - Information on operating conditions, assumptions, or limitations affecting precision and reliability. The evaluation will be based on the contracting authority’s overall assessment of the documented performance and supporting information related to precision and reliability in single-organism handling. Particular emphasis will be placed on, but not limited to, the following aspects: - Documented error rates - Ability to consistently perform successful single-organism transfers - Presence, functionality, and robustness of control and verification mechanisms - Documented success rate for single-organism handling and transfer Solutions demonstrating lower documented error rates, higher precision and reliability, and consistent performance in single-organism handling will receive a higher score. | B2
1.13 | The instrument should have the capacity to handling pipetting and sorting of live zebrafish (Danio rerio) eggs, embryos and larvae across a broad size range, up to 10 mm. Please provide the following information: - The maximum documented organism size (mm) for zebrafish (Danio rerio) eggs, embryos, and larvae that can be handled and pipetted reliably, specified as both maximum length (mm) and maximum thickness/diameter (mm). - Any technical prerequisites or limitations related to organism size (e.g. pipette tips, volume range, liquid flow characteristics, software settings). - Relevant documentation, test data, or references supporting the handling of larger organisms. The evaluation will be based on the contracting authority’s overall assessment of the documented capacity, reliability, and flexibility related to handling organisms across a broad size range. Particular emphasis will be placed on, but not limited to, the following aspects:: - Maximum documented organism size that can be handled reliably - Documented reliability when handling larger organisms - Identified technical limitations and dependencies - Quality, relevance, and credibility of the supporting documentation Solutions documenting reliable handling of larger organism sizes, with clearly described limitations and robust supporting documentation, will receive a higher score. | B3
1.14 | The instrument should support fluorescence-based detection and sorting of zebrafish (Danio rerio) eggs, embryos, and larvae. The solution should enable reliable identification and sorting based on fluorescence signal intensity and/or spectral characteristics, while maintaining gentle handling and high organism viability. Please provide the following information: A technical description of the fluorescence detection and sorting principles (e.g. excitation/emission wavelengths, optical configuration, detectors, sorting mechanism) - Supported fluorescence channels, wavelength ranges, and sensitivity (e.g. detection limits, signal-to-noise ratio) - Sorting performance data, including accuracy, precision, throughput (e.g. specimens per hour), and any limitations - Documentation or test data demonstrating sorting performance for zebrafish or comparable organisms - Description of how sorting parameters are configured (e.g. software-based gating, threshold settings) and ease of use The evaluation will be based on the contracting authority’s overall assessment of the documented fluorescence detection and sorting capability, with particular emphasis on: - Sensitivity and robustness of fluorescence detection - Sorting accuracy and reproducibility - Flexibility in configuration (e.g. channels, thresholds, software control) - Documented performance for relevant organism types and developmental stages - Throughput and operational efficiency - Impact on organism viability Solutions that offer high sensitivity, accurate and reproducible fluorescence-based sorting, flexible configuration, and well-documented performance with minimal impact on organism viability will receive a higher score. | B3
Training Requirements
1.15 | The supplier shall provide the necessary user training for approximately 2-3 users at the department. Training shall be conducted by agreement. The training shall ensure safe, proper, and correct use of the equipment. The supplier must expect that training may take place after handover. Describe the scope and content of the training. | A
Hygiene and Cleaning Requirements
1.16 | If there are any disinfectants or cleaning methods that cannot be used on the equipment, these should be described, together with an explanation of why the method is unsuitable | I
SERVICE, OPERATION AND MAINTENANCE
Warranty
2.1 | The warranty periode for the offered equipment shall be stated. | A
2.2 | An extended warranty period shall be offered as an optional feature and shall be priced in Appendix 3.0 Price Form- Guarantee option | A
2.3 | At the end of the warranty period, the equipment shall include the latest updated version of the software at no additional cost. | A
Requirements for service provisions
2.4 | The supplier shall provide prices for the following in Appendix 3.0 Price Form - servicetjenester - Call‑out fee: This shall cover all expenses and time associated with a round trip to and from the service location. - Hourly rate for service work: This shall apply to time spent performing service either at the contracting authority’s site or at the supplier’s premises. | A
Requirements for maintenance agreements
2.5 | An optional Level 1: Preventive Maintenance Agreement shall be offered. The maintenance agreement shall include: - Preventive maintenance in accordance with the manufacturer’s specifications, including labour, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - A response time guarantee, as specified in the supplier’s response to Requirement 2.11. Documentation of all work performed (service reports) shall be delivered to the contracting authority within 5 working days. Unless otherwise agreed, maintenance shall be carried out at the user’s site, with the time of service agreed in advance with the contracting authority. The agreement shall be priced in Appendix 3.0 Price Form - Maintenance Agreeement. | A
2.6 | An optional Level 2: Comprehensive Maintenance Agreement shall be offered. The maintenance agreement shall include: - Preventive maintenance in accordance with the manufacturer’s specifications, including labour, travel, and accommodation. - All wear parts, service kits, components, and batteries that the manufacturer recommends replacing as part of the maintenance. - All service costs, including labour and required parts. - A response time guarantee, as specified in the supplier’s response to Requirement 2.11. - All software updates, including new versions and any associated equipment updates. Documentation of all work performed (service reports) shall be delivered to the contracting authority within 5 working days. Unless otherwise agreed, maintenance shall be carried out at the user’s site, with the time of service agreed in advance with the contracting authority. The agreement shall be priced in Appendix 3.0 Price Form - Maintenance Agreeement. | A
Requirements for service and maintenance
2.7 | The delivery time for spare part should be as short as possible. Please specify the maximum delivery time from order to delivery at the contracting authority’s site, including any delivery from an external warehouse, expressed in number of calendar days. The shortest delivery time will receive the highest score. | B1
2.8 | A complete service manual in electronic format shall be provided with the delivery, if available. The service manual shall be in English or a Scandinavian language and shall remain accessible until the end of service. The customer shall receive all updates to the service manual, unsolicited and free of charge, for the entire duration of the service period as defined by the manufacture. | A
2.9 | Customer support by telephone shall be provided during regular working hours at no extra cost, until until the equipment reaches the end of service. Regular working hours are defined as weekdays from 09:00 to 16:00. Qualified technical personnel shall be available to answer questions. Communication with customer support shall be possible in a Scandinavian language or in English. The supplier shall provide a description of the customer support services available. | A
2.10 | A short response time is an advantage. The supplier shall state the response time for qualified technical personnel, by telephone or similar, when technical assistance is required during regular working hours. Regular working hours are defined as weekdays from 09:00–16:00. It is not permitted to submit reservations requiring the establishment of remote access in order to meet the response time. The response time shall be stated in minutes or, if necessary, hours, and is defined as the time from when the contracting authority reports a need for technical assistance until qualified technical personnel make contact to begin troubleshooting by telephone or similar. The shorter the stated response time, the higher the score. | B2
2.11 | The supplier should be able to provide a short on-site response time, specified in hours or calender days. Please specify the maximum on-site response time when urgent service is required after a fault has been reported during regular working hours for: - Maintenance Agreement Level 1: Preventive - Maintenance Agreement Level 2: Comprehensive - Call-outs without a maintenance agreement Regular working hours are defined as weekdays from 08:00 to 16:00. Evaluation will be based on the declared response times for each category. The shorter the stated on-site response time, the higher the score. | B2
2.12 | The supplier should offer a service solution that includes integrated remote monitoring of the instrument’s status, with the purpose of supporting efficient service, troubleshooting, and operational follow‑up. The service solution should enable both service personnel and users to monitor the instrument’s operating state, the status of ongoing runs, alerts in the event of errors or deviations, and other relevant operational information without physical presence at the instrument. Please provide the following information: - A description of the proposed service and remote monitoring solution. - Which functions are available to support service and operational follow‑up (e.g. diagnostics, alerts, condition monitoring, error logging). - Any limitations, prerequisites, or dependencies related to the use of the solution (e.g. network access, IT/security requirements, access control). The evaluation will be based on the contracting authority’s overall assessment of the documented functionality and suitability of the service and remote monitoring solution. Particular emphasis will be placed on the following aspects, which are not exhaustive: - Scope and functionality of the remote monitoring solution - Quality of operational overview and real‑time insight - Support for diagnostics, troubleshooting, and preventive maintenance - Usability and accessibility for both users and service personnel Solutions offering a high degree of functionality, good operational overview, effective service support, and user‑friendly access will receive a higher score. | B1
2.13 | The supplier shall provide a description of the procedure for performing preventive maintenance in accordance with the manufacturer’s instructions. Protocols for carrying out preventive maintenance, including information on maintenance frequency and the expected time required for each task, shall be attached. | A
2.14 | The supplier shall guarantee availability of spare parts for a minimum of 7 years after delivery of the equipment. | A
2.15 | Spare parts should be available for more than 7 years following delivery of the equipment. Please state the total number of years for which spare parts availability is guaranteed. The evaluation will be based on the overall assessment of the documented specifications and relevant characteristics related to long-term spare parts availability. Solutions that offer the longest guaranteed availability period will receive a higher score. | B2

## 3. Technical Specifications
TEKNISKE SPESIFIKASJONER
Art.nummer | Artikkelnavn | Leverandørens tilbud/modell: | Akustisk støy: (dB (A)) | Varmeavgivelse: Standby / drift (kW) | Krav til rommet: Temperatur (oC inkl peak-to-peak grenser) Luftfuktighet min/maks (% rH) | Strømforsyning: (stikkontakt, fast tilkobling, oppgi evt. strømkabellengde) Effektbehov (kW) Spenning (V) Strømstyrke (A) Krever separat kurs Krever separat jording | Spesielle krav til strømforsyning: Nødstrøm UPS | Datauttak posisjon og kabelføringsvei: Nettverkstilkopling (type tilkobling) Annen tilkobling Evt. krav til max kabellengde | Lyskilde Føringsvei | Krav til rommet: Temperatur (grader Celsius) Luftfuktighet min/maks (% RH) | Annet:
Art.nummer | Artikkelnavn | Leverandørens tilbud/modell: | Vanntilkobling: Type I eller type II posisjon og dimensjon eks. gjenger mm | Vanntilkobling: Byvann posisjon og dimensjon eks. gjenger mm | Krav til vannkvalitet: RO-vann | Krav til vannkvalitet: Tappvann/byvann | Krav til kjøling: eks. posisjon og dimensjon | Vannavløp: posisjon og dimensjon eks. gjenger | Avløp med spesielle rensekrav: posisjon og dimensjon eks. gips, fett, lo, olje | Lukket avløp: posisjon og dimensjon eks. dialyse | Sluk i gulv: JA/NEI | Formalintilkopling og andre væsker: eks. posisjon og dimensjon | Krav til redudans: | Krav til flowrate: (L/min) | Krav til flowrate: (L/min) | Forventet forbruk/døgn: ved normal drift og max drift | Annet
Art.nummer | Artikkelnavn | Leverandørens tilbud/modell: | Tilkoplinger til: Gass (trykkluft) Trykk (bar) Renhet (%) | Avtrekk: eks. luftutskiftinger ( /time) | Avtrekk: posisjon og dimensjon (Ø) | Avsug: posisjon og dimensjon eks. røyk, gass, anestesi (Ø) | Annet: