## 1. Guidance for use
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Project 1105601 UiO Livsvitenskap brukerutstyr
K937.02.02 MEA-system
Tenderer:
HOW TO FILL OUT THE DOCUMENT
Please note that there are several tabs to be filled out.
Columns and rows are not to be deleted.
The tenderer shall only fill out fields with a white background. Fields of other colors are not to be used or altered by the Tenderer.
The tenderer shall answer YES or NO to the question on whether or not a requirement is fulfilled. All rows must be answered.
The tenderer shall answer the defined requirements. The answer shall have sufficient information for the contractor to carry out a full assessment of the tender without the need to open attachments. If the tenderer needs to provide additional or more detailed information than this document allows for, it is allowed to refer to attachments. In case of reference to attachments, the tenderer shall provide the corresponding attachment nr and relevant pages.
Minimum requirements (A-criteria) are not related to an award criterion. Lack of fulfillment of minimum requirement can lead to rejecton of the tender, cf Tilbudsinvitasjon (Tender invitation).
B-criteria are related to an award criterion. The tenderers answer to B-criteria are the basis of the contractors evaluation and assessment. B-criteria are defined as B1, B2 or B3 criteria in the corresponding column for each criterion. The definitions indicate the assigned weight of the award criterion in the tender evaluation. Award criteria indicated as B3 are assigned most weight, while the B1 criteria are assigned the least weight.
Information criteria (I-criteria) are solely for informative purposes, and are not part of the evaluation.
| Valuta: | Lisens:
| NOK | Systemlisens
| SEK | Brukerlisens
| DKK | Utstyrslisens
| EUR | Analyselisens
| USD | Annet
GBP

## 2. Specification and response
SPECIFICATION OF TENDER AND SUPPLIER RESPONSE DOCUMENT
Criterion nr | Description of criterion | Criterion type | Supplementary information | Fulfilment of criterion Yes/No | Supplier response
1 | QUALITY AND FUNCTIONALITY CRITERIA
1.1 | The equipment shall be CE marked in accordance with all relevant directives and fulfil all relevant Norwegian laws, regulations, standards, requirements and norms for the equipment offered. Declaration of conformity shall be attached with the tender. | A
1.2 | The system shall, as a minimum, be suitable for its intended use as a research and educational tool for in vitro electrophysiology. The system is intended for performing extracellular electrophysiological recordings and controlled electrical stimulation in in vitro biological models at cellular and network levels under controlled laboratory conditions. It shall support research and educational activities in environments such as academic laboratories, core facilities, R&D laboratories and educational purposes for students and researchers. The primary purpose of the system is to enable quantitative investigation of extracellular electrophysiological function, including but not limited to: - Basic and applied neuroscience and cardiophysiology research - In vitro disease and mechanism‑of‑action studies - Pharmacological and toxicological investigations on electrogenic cell and tissue models - Development and validation of assays and methods for electrophysiological measurements and analysis. | A
1.3 | Provide the estimated delivery time from order reception to the equipment is delivered. If different parts of the equipment offered have different delivery times, this is to be stated here. | A
1.4 | Complete user manual(s) and operational documentation in English shall follow the delivery, also in Scandinavian language if possible. Preferably in electronic formats. The supplier shall, until end of service as given by the manufacturer, provide unsolicited electronic updates of the user manual(s), regardless of file format. | A
Equipment functionality
1.5 | The supplier shall describe the MEA technology offered (e.g. CMOS-based, passive/non-CMOS, or alternative approaches). The description shall include: - The underlying technology type and architecture - Key characteristics relevant to signal acquisition (e.g. on-chip vs external amplification, routing approach) - Any implications for system performance (e.g. electrode density, scalability, signal quality, flexibility) This information will not be evaluated, but is requested to provide an overview of the proposed solution. | I
1.6 | The system shall support extracellular electrophysiological recordings from the following biological models: – 2D cell cultures, including but not limited in neuronal and cardiac cells – 3D models, including neural-, cardiac- and retinal- organoid preparations, as well as brain slices The system shall be capable of accommodate these biological models for MEA‑based recordings under normal operating conditions for the intended use, as specified by the supplier. Compliance with this requirement shall be documented by describing the supported biological models and providing relevant technical specifications and/or application notes. | A
1.7 | The offered system shall support, at a minimum, the following primary sample types: - Brain organoids, ranging from approximately 500 µm to 2 mm in size - Cardiac organoids, ranging from approximately 200 µm to 500 µm in size Compliance with this requirement shall be documented. | A
1.8 | The system shall support extracellular electrophysiological recordings from electrogenic cells and provide stable signal performance during continuous recordings of at least 60 minutes under normal operating conditions for the intended use, as specified by the supplier. Signal stability shall be sufficient to enable reliable detection and quantitative analysis of electrophysiological events throughout the recording period. Compliance with this requirement shall be documented through relevant technical documentation. | A
1.9 | The system shall support simultaneous extracellular recording from a minimum of 64 electrodes per well, with all electrodes recorded concurrently during continuous acquisition under normal operating conditions for the intended use, as specified by the supplier Compliance with this requirement shall be documented through relevant technical documentation. | A
1.10 | The system shall support simultaneous extracellular recording from multi‑well MEA plates, with a minimum of 6 wells recorded concurrently under normal operating conditions. Compliance with this requirement shall be documented in the technical specifications, stating the minimum and maximum number of wells that can be recorded simultaneously. | A
1.11 | The system should support high electrode density, enabling extracellular electrophysiological recordings with single-cell or near single-cell spatial resolution. The supplier should describe the following: - The number of electrodes per well - The maximum number of electrodes within a selectable region of a well - Electrode size and inter-electrode spacing (pitch) - Any limitations affecting achievable spatial resolution - Documented capability for single-cell or near single-cell extracellular recordings The evaluation will emphasize: - Electrode density per well and within selectable regions - Electrode size and inter-electrode spacing as indicators of achievable spatial resolution - The extent and quality of documented capability for single-cell or near single-cell extracellular recordings, including any stated limitations Higher scores will be awarded to solutions demonstrating: - Higher electrode density - Smaller electrode size and tighter inter-electrode spacing - Clear, documented, and validated capability for single-cell or near single-cell extracellular recordings The evaluation will be based on the supplier’s response and submitted documentation. | B3
1.12 | The system should enable simultaneous recording from ≥1000 electrodes per well across a minimum of 6 wells. Please state: -The number of electrodes that can be recorded simultaneously per well -The maximum number of wells that can be recorded in parallel -The maximum number of electrodes that can be selected and recorded simultaneously per sample, where multiple samples may be present within a single well -Any limitations or dependencies affecting simultaneous recording capacity The evaluation will emphasize: - The number of electrodes that can be recorded simultaneously per well - The number of wells that can be recorded in parallel - The maximum number of electrodes that can be selected and recorded simultaneously per sample (if applicable) - The extent to which the specified performance is documented, validated, and free from limiting conditions Higher scores will be awarded to solutions demonstrating: - Simultaneous recording from a higher number of electrodes per sample and per well - Simultaneous recording across a greater number of wells - Clearly documented, substantiated, and robust performance with minimal limitations The evaluation will be based on the supplier’s response and submitted documentation. | B3
1.13 | The system should provide a clearly defined recording area with integrated microelectrodes, with a recording area larger than 2 × 2 mm. The supplier should specify: - Recording area dimensions (mm × mm) - Electrode density within the recording area - Any limitations related to the definition, coverage, or practical usability of the recording area The evaluation will place emphasis on: - The clarity and functional definition of the recording area, - The degree of microelectrode integration and electrode density, - The usable size of the recording area and any stated limitations. Solutions offering a more clearly defined and functionally integrated recording area, with better microelectrode integration and a larger recording area, will receive a higher score. The evaluation will be based on the users’ overall assessment of the documented specifications and relevant characteristics related to the recording area. The evaluation will be based on a comparative assessment of the submitted documentation against the stated evaluation criteria. | B3
1.14 | The system should provide sufficiently low noise levels to enable detection of small‑amplitude neuronal or cardiac signals, with a target (benchmark) noise level of < 6 µV RMS. Please provide: - RMS noise specifications, including the relevant measurement bandwidth(s) and measurement conditions, specified for each applicable electrogenic cell type and recording configuration. - Reference values or representative examples of signal‑to‑noise ratio (SNR) based on biological recordings, including a description of the biological model, recording conditions, and standard operating parameters used. - Any limitations or dependencies affecting achievable noise performance or biological signal quality. The evaluation will emphasize: - Documented RMS noise performance under well‑defined and relevant measurement conditions, - The quality and relevance of documented biological signal recordings, including reported SNR - The clarity and transparency of stated limitations or dependencies affecting noise and signal quality. Solutions that demonstrate lower RMS noise levels under clearly defined measurement conditions and provide well‑documented biological recordings showing high signal quality and favorable signal‑to‑noise ratios under realistic experimental conditions will receive a higher score. The evaluation will be based on a comparative assessment of the submitted documentation against the stated evaluation criteria. | B3
1.15 | The system shall support electrical stimulation with configurable stimulation amplitude and timing parameters enabling use in neuronal and cardiac network experiments. The system shall allow simultaneous electrical stimulation and extracellular recording under normal operating conditions for the intended use, as specified by the supplier. The system shall support adjustable stimulation amplitude, defined as the ability to configure stimulation signal strength within a specified range, resolution, and accuracy. This shall include: - Documented minimum and maximum stimulation amplitude values (e.g. in µV, mV, or µA, depending on the stimulation technology), - Specified amplitude resolution or step size, - Support for defined waveform polarity, including mono‑ and/or biphasic stimulation, with configurable phase amplitudes where applicable. The system shall support configurable stimulation timing parameters, including at a minimum adjustable pulse duration, stimulation frequency and inter-pulse interval. Compliance with this requirement shall be demonstrated through technical specifications describing the stimulation capabilities, including supported stimulation modes, configurable parameters, and relevant performance characteristics. | A
1.16 | The system should offer high stimulation capacity and flexibility suitable for advanced neuronal network experiments. The system should: - Support simultaneous electrical stimulation of a high number of electrodes per well (minimum 26), with flexible and user‑selectable electrode configurations for each well, - Support parallel electrical stimulation across user‑selectable wells or across all wells of a compatible MEA plate, - Enable high‑resolution electrical stimulation allowing targeted activation of individual neurons or small neuronal populations. The evaluation will emphasize - The number of electrodes that can be stimulated simultaneously per well, - The degree of flexibility in electrode selection and configuration, - The ability to perform parallel stimulation across multiple or all wells, - The documented spatial resolution and capability for targeted electrical stimulation, - Any stated technical limitations or dependencies affecting stimulation performance. Solutions offering a higher number of simultaneously stimulated electrodes per well, greater flexibility in electrode and well selection, and documented capability for high-resolution, targeted stimulation will receive a higher score. The evaluation will be based on documented specifications and verifiable characteristics described in the supplier’s submission. | B3
1.17 | The system should support optical stimulation, including functionality suitable for optogenetic applications. The supplier should describe the available optical stimulation functionality and provide relevant supporting documentation. Any limitations or dependencies should be clearly stated. The evaluation will emphasize: - The documented optical stimulation capabilities and technical characteristics, - The suitability of the optical stimulation functionality for intended applications, including optogenetics, - The degree of integration and practical usability of the optical stimulation solution, - The clarity and completeness of the provided documentation and stated limitations. Solutions that offer more advanced, well‑integrated, and clearly documented optical stimulation functionality, with demonstrated suitability for intended applications and good practical usability, will receive a higher score. The evaluation will be based on a comparative assessment of the submitted documentation against the stated evaluation criteria. | B3
1.18 | The system shall provide precise control and stable maintenance of environmental conditions for electrophysiological recordings. The system shall maintain the following conditions in the sample environment before and throughout electrophysiological recording sessions under normal operating conditions for the intended use, as specified by the supplier: - Temperature: 37 °C ± 0.5 °C, and - CO₂ concentration. 5 % ± 2 %, The supplier shall document compliance by providing technical specifications describing temperature and CO₂ control performance, including stated tolerances and operating conditions. | A
1.19 | The system shall support acute, short‑term, and long‑term extracellular electrophysiological recordings under normal operating conditions for the intended use, as specified by the supplier, defined as follows: – Acute recordings are defined as recordings performed on freshly prepared biological samples, typically with durations ranging from minutes to several hours. - Short‑ and long‑term recordings are defined as continuous or repeated recordings over extended culture periods. Short‑term recordings refer to durations of several days up to two weeks, while long‑term recordings refer to durations of several weeks to months. Compliance with this requirement shall be documented by describing the supported recording modes, achievable recording durations, and any relevant limitations. | A
1.20 | The system shall be delivered with dedicated software that supports real-time data acquisition, visualization, and electrophysiological data analysis. The software shall be fully operable offline and shall not require an internet connection for normal operation. The software shall support export of raw data in both uncompressed and compressed formats using standard, non-proprietary file formats (e.g., .brw, CSV, HDF5, or equivalent). The software shall, as a minimum, include the following functionalities: - Spike detection and electrophysiological activity characterization, including firing rate, spike amplitude, and data filtering - Analysis of bursting activity and synchronized network behavior - Generation of customizable analysis reports including selected recordings and analysis outputs such as raw traces, activity maps, space-time heatmaps, waveform plots, raster plots, trend charts, histograms, box plots, and spike and field potential metrics The supplier shall provide documentation and tutorial resources supporting external and offline analysis workflows using exported data and metrics. Core functionalities shall not require recurring license fees. Any optional licensed modules shall be clearly specified. Compliance with all elements of this requirement shall be documented. Relevant technical documentation shall be attached to the tender response. | A
1.21 | The solution should offer a comprehensive set of advanced analysis functionalities, including, but not limited to: - axonal signal tracking and propagation analysis - spike sorting - network-level analyses - AI-based spike event validation for automated quality control - batch analysis across multiple recording files. The software should demonstrate a high level of maturity, including robustness and stability of the analysis functionalities, documented validation of methods, and evidence of practical use in relevant applications and provide well-structured and comprehensive documentation covering both functionality and usage of the analysis tools. The software should also demonstrate high usability, including ease of use, efficient execution of analysis tasks, and an intuitive and well-structured user interface and, as well as provide programmable API access enabling custom workflows, automation, and integration with external tools. The licensing model should be based on a perpetual license and should include software upgrades at no additional cost. Higher scores will be awarded to solutions that demonstrate a broad and advanced set of well-integrated analysis functionalities, mature, robust, and well-documented software with validated methods and proven use, as well as high usability, including efficient workflows and an intuitive user interface. Furthermore, solutions offering strong integration capabilities through flexible API access, and favorable licensing terms, including perpetual licensing and included upgrades, will achieve a higher score. | B3
1.22 | The system should support assessment of cell and 3D structure attachment and viability, including applications such as organoid preparations and other 3D biological models. The supplier should describe the available methods and functionality for assessing attachment and viability, including any limitations, and provide supporting documentation The evaluation will emphasize: - The documented methods and functionality available for assessing cell and 3D structure attachment and viability, - The relevance and applicability of the described functionality to organoid and similar 3D models, - The clarity and completeness of the provided documentation, including stated limitations. Solutions that provide clear and well‑documented support for assessment of cell and 3D structure attachment and viability, with demonstrated applicability to relevant biological models, will receive a higher score. The evaluation will be based on a comparative assessment of the supplier’s response and submitted documentation against the stated evaluation criteria. | B3
1.23 | The system should provide a solution that ensures secure, stable, and consistent sample positioning, cell attachment and electrical contact of organoids with the electrodes throughout the experimental duration, without causing observable tissue damage. The solution may consist of dedicated holder systems or alternative approaches, such as liquid confinement systems, surface coatings, or other equivalent positioning strategies. All proposed approaches will be considered, provided they fulfil the functional requirements outlined below. The supplier should describe: - The design and operating principle of the proposed solution - How stable sample positioning, and electrical contact are achieved and maintained over time - Compatibility with different organoid types, sizes, and experimental conditions - Practical usability, including handling, manipulation, and integration into experimental workflows - Any limitations, dependencies, and conditions required for optimal performance Please attach relevant supporting documentation where applicable. Both holder-based solutions and alternative approaches will be evaluated equally, provided they meet the stated functional requirements. The evaluation will place emphasis on: - Documented ability to ensure stable and consistent positioning, attachment and electrical contact of organoids with the electrodes over time without causing observable tissue damage - Compatibility with a range of organoid types and sizes - Practical usability, including ease of handling and workflow integration - Scope of limitations and dependencies - Degree of flexibility and applicability across use cases Solutions that demonstrate robust and well-documented performance, broad compatibility and minimal limitations or dependencies will achieve a higher score. The final evaluation will be based on the contracting authority’s overall assessment of the supplier’s response and submitted documentation. | B3
1.24 | All components, consumables, or accessories required to ensure stable positioning and contact of organoids during normal operation of the system shall be included in the supplier’s offer. For solutions relying on physical holders, the following minimum quantities shall be included per holder type: - A minimum of fourteen (14) units for systems based on 6-well configurations - A minimum of twenty-five (25) units for systems based on 24-well configurations Alternatively, the supplier may propose a different number of units per holder type, provided that the proposed quantity is clearly demonstrated to be sufficient to support initial experimental use of the system. Such deviations shall be justified and documented. | A
1.25 | The supplier should provide a comprehensive and structured description of the solution(s) used to ensure positioning of organoids and stable electrical coupling to the electrodes, as referenced in criterion 1.23. For holder-based solution, the supplier should provide an overview of holder types included in the solution. For non-holder-based solutions, the supplier should describe: - The positioning and electrical coupling principle applied - Any associated components and consumables For all solutions, the supplier should describe: - Supported organoid types, sizes, and experimental conditions - Relevant limitations, dependencies, and compatibility constraints - Requirements related to equipment, consumables, and operating conditions - Practical usability, including setup, handling and integration into experimental workflows - Whether the solution supports: - Multiple alternative coupling methods - Acute recordings - Long term recordings (including where different approaches may be used) - Compatibility with third-party equipment (non-supplier-specific products) The evaluation will place emphasis on: - Clarity, structure, and completeness of the description of the proposed solution - The extent to which compatibility with different organoid types, sizes, and experimental conditions is clearly described and documented. - Practical usability as described in the documentation, including ease of implementation, handling, setup, and integration into experimental workflows - The extent of limitations, constraints, or dependencies associated with the solution, including requirements for specific equipment, consumables, or operating conditions. - The extent to which the supplier provides multiple validated coupling solutions, including solutions specifically suited for acute recordings, and solutions compatible with non-in-house products. Solutions that demonstrate broad applicability, strong practical usability, flexibility through multiple documented and validated approaches, particularly those supporting acute recordings, long-term recordings, and compatibility with non-in-house or third-party equipment, will achieve a higher score. Relevant technical specifications, product descriptions, or other supporting documentation shall be attached to the tender response. | B1
1.26 | The offer shall include a sufficient number and mix of compatible MEA plates to enable initial experimental use of the system across all supported configurations and application types. The supply shall, as a minimum: - include a total of at least 10 MEA plates, and - If the system offers both 6-well and 24-well MEA plate formats, the offer shall include at least 5 6-well MEA plates and 3 24-well MEA plates, and - ensure that all supported application types (e.g., 2D planar cultures and 3D cultures and/or organoid applications, where applicable) can be used for initial experiments. All MEA plates shall be fully compatible with the offered system and included in the delivery. The supplier shall clearly specify: - total number of MEA plates included - distribution across well formats - intended application types supported by each plate type Compliance with this requirement shall be documented. Relevant supporting documentation shall be attached. | A
1.27 | The MEA plates and associated solutions should be compatible with optical and/or live‑cell imaging. Compatibility may be achieved through plate design, or through additional accessories or solutions provided by the supplier. The supplier should describe: - How compatibility with optical and/or live‑cell imaging is supported (e.g. plate material, transparency, geometry, or supporting devices/solutions) - Any additional equipment or configurations required - Any limitations or dependencies If any additional supporting devices or accessories are required to achieve compatibility, these shall be clearly specified and included in the quotation, with stated pricing. Relevant supporting documentation shall be provided. The evaluation will be based on: - The documented compatibility of the MEA plates with optical and/or live‑cell imaging - The practicality and usability of the proposed solution (including any supporting devices or setups) - The level of flexibility offered for different experimental workflows - The clarity and completeness of the documentation, including stated limitations Solutions that provide well‑designed MEA plates and/or practical supporting solutions enabling effective optical and/or live‑cell imaging, with minimal constraints and good usability, will receive a higher score. | B3
1.28 | MEA plates shall have a documented shelf life or expected usable lifetime of minimum 6 months from the date of delivery. The supplier shall document either: - a defined shelf life/expiry date (from date of delivery), or - an expected usable lifetime under normal storage and usage conditions. The documentation shall, as a minimum, include: - the stated shelf life/expiry period and/or expected usable lifetime - specified storage and handling conditions that may affect shelf life or usable lifetime - criteria defining when the plates shall be considered expired or require replacement - any applicable warranty terms, including conditions for replacement in case of failure during first use If an expiry date is applied to the product, this shall be clearly indicated on the plates and/or packaging. If no formal expiry date is defined, the supplier shall provide substantiated documentation describing the typical durability and performance stability of the plates over time. Supporting documentation shall be provided as part of the tender. | A
Training
1.29 | Necessary training of super users for approximately 3 pax shall be offered. The training is to take place in agreement with the contractor. The purpose of the training is to ensure secure and correct us of the equipment for the intended use. The supplier must expect that the training will take place after delivery and hand-over of the equipment. Describe the scope and content of the training offered. | A
1.30 | The supplier should offer structured training and post‑installation support beyond the initial system installation. This may include: – Structured, multi‑session training programs – Access to online training resources, such as videos, protocols, and software updates – Access to dedicated application scientists or equivalent resources for post‑installation support The evaluation will emphasize: – The scope and structure of the training provided – The accessibility of training resources and support – The relevance of the training and support for practical system use Solutions that provide comprehensive, structured, and well‑documented training and post‑installation support, with good accessibility and relevance for practical use, will receive a higher score. The evaluation will be based on a comparative assessment of the supplier’s submitted documentation against the stated evaluation criteria. | B2
Cleaning and hygiene
1.31 | In case of incompatible disinfection agents and/or methods that cannot be used for cleaning of the equipment, this shall be described here, accompanied by a rationale for the incompatibility. | I
1.32 | Describe recommended procedure(s) for cleaning, disinfection and/or sterilization of the equipment offered. | A
1.33 | The equipment should be designed to support safe, practical, and efficient laboratory use, including handling of MEA plates and any associated sample holders. The supplier should describe: - Recommended procedures for routine cleaning and general maintenance of system surfaces and components exposed during normal use - Compatibility with standard laboratory cleaning agents and disinfectants - Handling, cleaning, and reuse (if applicable) of MEA plates, sample holders, and any plate adapters or fixtures - Any specific requirements or limitations related to contamination control - The expected level of user effort required for routine cleaning, plate handling, and maintenance The evaluation will consider: - Ease of maintaining a clean and contamination-controlled working environment - Practicality of handling MEA plates and sample holders in day-to-day workflows - Clarity and completeness of the documented procedures - Overall usability and efficiency in typical laboratory use The evaluation will be based on the users’ overall assessment of the supplier’s response. | B1
ICT
1.34 | All mandatory hardware, middleware, software and security updates are to be provided and installed without cost until end of service (updates are to be interpreted as necessary improvements and bugfixes). | A
2 | SERVICE. OPERATION AND MAINTENANCE CRITERIA
Warranty
2.1 | Provide warranty for the offered equipment. | A
2.2 | Extended warranty period shall be offered as an option and priced in Appendix 3.0 – Pricing Schedule. | A
2.3 | Upon expiry of the warranty period, the equipment shall have the latest software version, without further costs. | A
2.4 | During the warranty period, the supplier shall provide service equivalent to Maintenance Agreement Level 2 (Comprehensive). Criteria 2.7 describes what Maintenance Agreement Level 2 includes. Periodic maintenance shall be carried out in accordance with the manufacturer’s recommendations. The final periodic maintenance shall be performed as the warranty period approaches its expiration. For example, a warranty period of one year will entail one instance of periodic maintenance if the manufacturer recommends annual maintenance. Documentation of completed work (service report) shall be submitted to the contracting authority within five (5) working days. Unless otherwise agreed, maintenance shall be performed at the user’s premises. The supplier shall agree on a suitable time for execution with the contracting authority in advance. | A
Service requirements
2.5 | The supplier shall provide prices for the following items in Appendix 3.0 Pricing schedule: -Attendance fee (shall include all costs and time consumption pr roundtrip to service address -Hourly rate for service (including used time at the Contractor or supplier) | A
Maintenance agreement
2.6 | Option for Level 1 Preventive maintenance agreement shall be offered. The maintenance agreement shall include: -Preventive maintenance in accordance with the manufacturers specifications, including all wear parts/service kits/components/batteries that the manufacturer recommends to be swapped during maintenance -Service report (documentation of performed maintenance) shall be supplied to the Contractor within 5 working days Unless agreed otherwise, the maintenance is to be performed at the Contractor address. The supplier must agree in advance on the time for the performance of the service with the Contractor. The agreement shall be priced in Appendix 3.0 Pricing schedule. | A
2.7 | Option for Level 2 - Full maintenance agreement shall be offered. The maintenance agreement shall include: -Preventive maintenance in accordance with the manufacturers specifications, including all wear parts/service kits/components/batteries that the manufacturer recommends to be swapped during maintenance -All service costs, including costs related to work and all consumables -Service hour guarantee, cf suppliers answer to criteria 2.13 on service level agreement -Response time guarantee, cf suppliers answer to criteria 2.11 -Regular security updates of antivirus protection and operating system -All software updates, including new versions and corresponding hard- and middleware updates -Service report (documentation of performed maintenance) shall be supplied to the Contractor within 5 working days. Unless agreed otherwise, the maintenance is to be performed at the Contractors address. The supplier must agree in advance on the time for the performance of the service with the Contractor. The agreement shall be priced in Appendix 3.0 Pricing schedule. | A
Service and maintenance
2.8 | Short delivery time for spare parts is preferred. Provide the maximum expected delivery time from confirmed order to delivery to Contractor, including deliveries from external warehouses. (1 day= 24 hrs). The delivery time specified by the supplier will be evaluated. Shorter delivery times will be considered an advantage and will result in a higher score in the evaluation. | B1
2.9 | Complete electronic service manual shall accompany the delivery, if available. The manual shall be in English and/or a Scandinavian language. The Contractor shall, until end of service as stated by the manufacturer, be provided with unsolicited access to updates of the service manual, without additional costs. | A
2.10 | The tender shall include customer service by telephone within ordinary working hours without further costs. Qualified technical personnel shall be available to answer customer questions. Communication with the customer support shall be conducted using a Scandinavian and/or English language. Ordinary working hours are defined as weekdays 8 AM - 4 PM. Free customer support shall be included until end of service. Describe available customer support. | A
2.11 | Short response time is preferred. Provide expected response time for qualified technical personnel, per telephone or equivalent, in case of need for technical assistance within ordinary working hours. Ordinary working hours are defined as weekdays 8 AM-4PM. Is is not allowed to demand establishment of remote access in order to comply with the provided response time. Response time is to be provided in minutes, or hours, and is defined as the time from the Contractor reports a need for technical assistance until qualified technical personnel makes contact to start troubleshooting by telephone or equivalent. | B1
2.12 | The supplier should have a service organization with short response time for troubleshooting of acute matters. Provide information about the maximum response time from reception of a trouble call until a qualified person is on site for acute service within ordinary working hours for the following: -Level 1 Preventive maintenance agreement -Level 2 Full maintenance agreement -Service call without maintenance agreement Ordinary working hours are defined as weekdays 8AM-4PM. | B1
2.13 | Long service hours guarantee is preferred. Provide information on the service hour guarantee within ordinary working hours, over a 6 month period for: -Level 1 preventive maintenance agreement -Level 2 full maintenance agreement -During the warranty period Ordinary working hours are defined as weekdays 8AM-4PM. | B1
2.14 | Provide a brief description of the manufacturers recommendation for preventive maintenance. Protocols for preventive maintenance, as well as information on frequency of maintenance and expected time consumption is to be attached. | A
2.15 | In case any special tools, calibration tools, software, passwords, service key or such is required in order to perform technical troubleshooting, maintenance and/or acute service, this is to be stated here and any associated costs are to be listed in Appendix 3.0 Pricing Schedule. | A
2.16 | Describe which parts that are expected to be swapped during the anticipated life time of the equipment, and the swapping frequency. | A
2.17 | The supplier shall guarantee access to spare parts for a minimum of 7 years after delivery of the equipment. | A
2.18 | Access to spare parts beyond 7 years after delivery will be awarded. Provide the total number of years the supplier guarantees access to spare parts. | B1

## 3. Technical specifications
TECHNICAL SPECIFICATIONS
Statsbygg art. nr | Article name | Tenderers offer/model | Acoustic noise: (dB (A)) | Heat emission: Standby / operation (kW) | Room requirements: Temperature (oC incl peak-to-peak limits) Relative humidity min/max (% rH) | Power supply: (power plug, fixed plug, length of power cord) Effect (kW) Voltage (V) Ampere (A) Require dedicated circuit Require dedicated earthing | Specific electricity requirements: Emergency power supply UPS | Ethernet: Socket position, cable type, cable routing, plug type, other, capacity Any max cable length | Lighting source Cable routing | Other:
029.05.003:1 | MEA-system
Statsbygg art. nr | Article name | Tenderers offer/model | Water supply: Type I, II or III Outlet position, dimensions, threads (mm) | Water supply: City water Outlet position, dimensions, threads (mm) | Water quality parameters: Type I, II or III | Water quality parameters: City water | Cooling/chiller requirements: e.g. positions, threads (mm) | Water drain: Position, dimension s (e.g. threads (mm)) | Drains with requirements for treatment (e.g. plaster, fat, lint, oil, biologicals,chemicals): Positioning, dimensions | Closed drain: Positioning, dimensions (mm) | Floor drain(s): YES/NO Positioning | Redundancy requirements: Type | Flow rate: -L/h and L/min | Expected 24h consumption during normal and max usage | Other:
029.05.003:1 | MEA-system
Statsbygg art. nr | Article name | Tenderers offer/model | Couplings to: Gases (incl compressed air) Pressure (bar) Purity (%) | Exhaust: e.g air changes/h, dedicated or shared | Exhaust: Positioning, dimensions (Ø) | Process exhaust: (e.g. fume, gas, chemicals) type Positioning and dimensions (Ø) | Other:
029.05.003:1 | MEA-system