## Behov SSA-K
Anskaffelse: 26178 Analyseinstrument til lab
Nr | Behov | Dekkes (S, T, N) | Tilbyders Kommentar
1.1 | General
1.1.1 | The system should be a hybrid multi-mode plate reading and automated cell imaging instrument for cell-based assays, kinetics, and long-term live imaging including automated analysis within the same platform
1.1.2 | The system is able to perform several reading methods including well-area scanning, kinetic, and end point
1.1.3 | The system should be documented to be suitable for both biochemical assays and advanced cell-based imaging assays
1.1.4 | The system is able to integrate environmental control (temperature/CO2/O2) for long-term live-cell assays
1.1.5 | The system is compatible with 6, 24, 96, and 384-well microplates
1.1.6 | The system is compatible with microscope slides and T25 cell culture flasks
1.1.7 | The system should have the capacity of at least 8 microplates for automatic live cell imaging, or be ready for upgrading including addition of attached automated incubator with at least 8 plate capacity.
1.1.8 | The system is able to hold data analysis software for quantitative and qualitative analysis within the same instrument platform
1.1.9 | The system is capable of performing fluorescence, absorbance, and luminescence microplate reading
1.1.10 | The system is able to provide imaging modes including fluorescence, brightfield, color brightfield, phase contrast, and high‑contrast brightfield.
1.1.11 | The system should support integrated dual reagent injector, enabeling fast, kinetic response tests.
1.1.12 | The system should have a cooling module for rapid cooling after performance.
1.1.13 | The platform should be modular and upgradable, including expansion of imaging functionality, installing new objects and filter cubes. The vendor should describe available upgrades which can be performed on site at the customer.
1.1.14
1.2 | Imaging
1.2.1 | Imaging methods should include single image, time lapse, and z-stacking.
1.2.2 | The instrument should be equipped with at least 4 fluoresence channels.
1.2.3 | The instrument should have a brightfield channel.
1.2.4 | The instrument should be equipped with objectives: 4x, 10x, 20x, 40x.
The vendor should be able to deliver a 40X extended plan apochromat objective with minimum 0,95 NA
1.2.5 | The instrument should offer 1.25x to 60x magnification with inverted microscope system.
1.2.6 | The instrument should have a capacity of at least 6 objectives.
1.2.7 | The vendor is able to supply a 16‑bit depth scientific-grade area-scan sensor (sCMOS or equivalent) camera.
1.2.8 | The platform should support brightfield, high contrast brightfield, phase contrast og fluorescense imaging within the same instrument without the need of external modules.
1.2.9 | The system should have integrated inverted widefield fluorescense microscopy
1.2.10 | The instrument should have autofocus and auto-exposure.
1.2.11 | The instrument should provide user-trained autofocus.
1.2.12 | The instrument should have a dedicated laser-based autofocus for fast and reproducible measurements, including support for thicker biological samples and 3D-models.
1.2.13 | The system should have an imaging filter cube capacity of up to four simultaneous user-replaceable fluorescence cubes, in addition to brightfield channel.
1.2.14 | The vendor is able to provide filter cubes for at least the following fluorophores (or equivalent spectral coverage): DAPI (or Hoechst), FITC/GFP, Cy3/Alexa 555 (or PE), Texas Red/TRITC (or Alexa 594), and Cy5 (or Alexa 647).
1.2.15 | The vendor s is able to provide additional filter cubes for CFP (and CFP–YFP FRET), YFP, Cy7/Alexa 750, and PI/Acridine Orange as needed.
1.2.16 | The system is able to switch between imaging channels automatically via the software control.
1.3 | Performance
1.3.1 | The vendor should provide camera field of view with 4x-objective, in addition to estimated imaging times for one whole 96-well plate.
1.3.2 | The system should be able to include detection modes: fluorescence intensity, absorbance; luminescence.
1.3.3 | The instrument should be able to perform automated imaging in brightfield and phase contrast within the same analysis protocol, without any user intervention.
1.3.4 | The vendor is able to demonstrate compatibility of the system with 3D-models, organoids and long-term live-cell imaging.
1.3.5 | CO2 control should be in the range 0 - 20%
1.4 | Software/analysis
1.4.1 | The system should include an image analysis module for neurite outgrowth
1.4.2 | The system should include an image analysis module for spot counting.
1.4.3 | The system should include an image analysis module for automated region of interest (ROI) identification.
1.4.4 | The software handles plate reading, automated imaging, image analysis, quantitative data analysis and kinetic analysis without the need of any third-party software.
1.4.5 | The system software supports cell counting, confluence analyses, viability-analyses, translocation-analyses, 3D-models, organoid imaging, whole-well montage and phenotypic analyses.
1.5 | Computer
1.5.1 | The vendor should provide minimum and recommended computer specifications and approve the customer-supplied workstation prior to shipment. Alternatively, the vendor may supply a preconfigured, validated workstation.
1.5.2 | The vendor should install and configure all required software, drivers, and licenses.
1.5.3 | The software control and analysis modules should not come with any licence costs.
1.6 | Assays/detection modes
1.6.1 | The system should be capable of performing live cell imaging.
1.6.2 | An absorbance reading module should be available as an option.
1.6.3 | The system should be equipped with a fluoresence reading module.
1.6.4 | A luminescence reading module should be available as an option.
1.7 | Installation
1.7.1 | The vendor should perform on-site delivery, installation, and setup of the complete system, including all accessories and software. The freight and installation is included in the price.
1.7.2 | The vendor should provide manuals, safety data (if any), and calibration certificates at installation.
1.7.3 | The vendor should offer on‑site training for at least 2 primary users.
1.8 | Service
1.8.1 | The vendor is able to provide on-site preventive maintenance service annually.
1.8.2 | The vendor should be available to assist with technical issues (via telephone, mail, or on site) during normal business hours and provide timely guidance to restore operation.
1.8.3 | The vendor should have a service- and support organization established within Norway, with trained service technicians.
1.8.4 | The vendor should have a short response time.
1.9 | Support
1.9.1 | The vendor is able to provide unlimited user-support free of charge.