## Information
Appendices 1 and 2
ANSK-26-0135: Ultrafast spectroscopy system
Information to the supplier:
The requirements in the table of requirements form the basis for the Supplier’s description of its solution specification cf. Appendix 1 and 2. The Supplier must give an extensive description of all requirements so that the Parties will have a common understanding of what is to be delivered, or which can give the Buyer alternative perspectives to the requirements expressed in the Buyer’s requirements specification.
The different requirements have different priority.
Priority A = Absolute requirements that must be fulfilled. Tenders that do not fulfill all the absolute requirements will be rejected.
Priority B = Important requirements that should be fulfilled, but it is not an absolute requirement. Answers will be of great importance for the evaluation of the tender. The requirements are weighted higher than C-requirements.
When filling out the table of requirements the Supplier shall mark an (X) in the column for yes, partly or no under “the Supplier’s fulfilment”. Reference to relevant documentation and any additional comments shall be made in the column “Comment”. It is reminded that it is the duty of the Supplier to document that it fulfils the set requirements. Any missing documentation will result in the requirement being considered as not fulfilled.
Document contents:
This document consists of Appendices 1 and 2. Make sure that all the tabs are filled out and answered according to the buyers requirments.

## Appendix 1_2
Appendix 1_2
ANSK-26-0135: Ultrafast spectroscopy system
Req. | Priority: | Description: | The supplieres fulfilment: |  |  | Response/comments:
|  |  | Yes | Partly | No
General requirements | [5 empty] | General requirements
1 | A | Deliverables: The offer must contain all elements needed for the equipment/instrument to function as intended. The offer shall include all deliverables, such as: - The delivery of all items included in the offer, after formal order from the buyer - Documentation in English for the system/equipment - Software, which should be the latest updated version at the point of delivery - List of spare parts with prices - Provision of all documents pertaining to the equipment to include, but not limited to, maintenance manuals, operation manuals and commissioning certificates. Quote the price in Appendix 3, Table 1.
2 | A | Warranty: The system/equipment shall be offered with a minimum of 12 months of warranty. The warranty period shall begin after the Site Acceptance Test (SAT). Content and coverage of warranty must be specified. Warranty period starts when the delivery has been accepted through a Site Acceptance Test (SAT). To make sure that both parties accept the delivery, there should be a mutual signing of the attached approval protocol.
3 | A | Training: Physical training for use and basic maintenance must be included in the offer. The offer must include on-site training involving a minimum of 1 of the personnel from the supplier, and a minimum of 3 from the buyer. In addition, both professional and technical staff from the buyer must have the opportunity to be present at the time/date of delivery and signing of the approval protocol.
4 | B | Training: We ask the offeror to specify: - The content/subjects included in the training - How the training will take place and the length of the training - Number of people from the buyer who can participate in the training (minimum 3) - Type and locations of other additional courses/training offered by the supplier (if any)
5 | A | Support: E-mail and telephone support shall be included in the offer. If remote support is available for the offered equipment/syetem, this shall also be included. Other available support methods shall be mentioned in the supplier's response. If there are conditions for such a support, this should be described here.
6 | B | Response time: Normal response time should be described by the supplier (e.g. hours, days) between a fault/error being reported in writing by the buyer and the supplier offering a solution.
7 | A | Claims within the warranty period: If problems occur during the warranty period, the supplier shall work continually to resolve the problem.
8 | B | If the supplier is unable to repair the problem within the response time given, the supplier should offer loan parts or loan of equipment until the problem is resolved. Conditions attached to such a solution should be described in the response column in appendix 2.
9 | A | Electricity/power: The equipment shall accept Norwegian power supply (240 V, 50 Hz or 380 V/400 V/50 Hz). The system should be simple to upgrade with respect to forthcoming hardware and software. Please describe the offered solution for electrical and other technical preferences. To expand on your answer, please use our appendix 4 in the tender document.
10 | A/B | Service: The Supplier shall submit offers for the following maintenance agreements applicable during the warranty period (first year): - Full service agreement (including travel costs, spare parts, etc.) - Maintenance agreement (excluding travel costs, spare parts, etc.) The Contracting Authority will include the prices for both of the above-mentioned maintenance agreements in the price evaluation. Please specify what is included in each maintenance agreement and quote the corresponding prices in Appendix 3, Table 2.
11 | A | All costs from the supplier concerning the offer, installation, travelling costs etc are to be included in the offer given in this tender. This means all costs for the supplier up until the SAT and formal takeover. Quote the price in Appendix 3, Table 1.
12 | A | Delivery and signing of approval protocol: Upon delivery of the equipment, an acceptance protocol shall be signed by both the Supplier and the Contracting Authority. Any deviations from the tender or the Contract shall be documented, and a timetable for corrective actions shall be agreed. The warranty period shall not commence until any material deviations have been corrected and the equipment has been formally accepted by the Contracting Authority.
13 | B | Demonstration The supplier should be able to present the offered solution for a demonstration if the buyer should choose to have the system demonstrated. The buyer should have the right to a demonstration but shall not be obligated to do so. The demonstration should be presented at a location where the offered system is being used. Please describe how this could take place. Please refer to sites where the offered system is being used.
14 | B | Delivery time: Enter the expected delivery time, in weeks, for the offered product/instrument from the date of the purchase order (PO). Shorter delivery times will be favoured.
Technical requirements |  |  | Yes | Partly | No
15 | A | Integrated system for ultrafast transient absorption spectroscopy, comprising the following modules: · (i) femtosecond laser · (ii) OPA · (iii) transient absorption spectrometer · (iv) extension for long lifetime measurements · (v) software. Technical requirements for each module are set out below.
16 | A | Femtosecond laser with the following characteristics: · (i) pulse width: <= 190 fs · (ii) pulse energy: ≥ 300 uJ (micro Joules) · (iii) average output power > 5 W · (iv) tuneable repetition rate from single shot to at least 100 kHz · (v) pulse energy stability <0.5 % RMS over a continuous 24-hour period after warm-up
17 | B | Please specify pulse width of laser (shorter is better)
18 | B | Please specify max pulse energy of laser (higher is better)
19 | B | Please specify average output power of laser (higher is better)
20 | B | Please specify type and fundamental wavelength of laser.
21 | B | Please specify RMS pulse-to-pulse energy stability over a continuous 24-hour period after warm-up. (lower is better)
22 | A | Compatible OPA ( Optical Parametric Amplifier) with the following characteristics: · (i) spectral range: 320 – 2600 nm, · (ii) beam diameter < 8 mm, · (iii) pulse-to-pulse power variation < 2 %, · (iv) automated, hands-off switching between wavelengths, · (v) single port output for all wavelengths
23 | B | Please specify full spectral range of OPA, stating any gaps in coverage (wider is better, fewer gaps is better)
24 | B | Please specify beam diameter after OPA (smaller is better)
25 | B | Please specify pulse-to-pulse power variation from OPA (lower is better)
26 | B | Please provide OPA output power (or pulse energy) as a function of wavelength across the full tuning range.
27 | B | Please specify M2 beam quality after OPA (lower is better)
28 | B | Please specify pump arrival time variation at the sample across the full tuning range without user adjustment. (lower is better)
29 | A | Transient Absorption Spectrometer with the following characteristics: · (i) instrument response function (IRF) < 200 fs · (ii) probe diameter at sample: < 50 um FWHM · (iii) sample chamber can accommodate samples of at least 20 cm x 20 cm · (iv) probe range: 350 – 1600 nm or wider · (v) automated neutral density filters for pump and probe, · (vi) delay line and detector characteristics as specified below · (vii) dedicated optics (lens, ND filter, iris, crystal) for white light continuum in different probe wavelength ranges (UV, VIS, NIR)
30 | B | Please specify full width at half maximum (FWHM) of the instrument response function (IRF), measured as the pump–probe cross-correlation under the supplier’s standard operating conditions. (lower is better)
31 | B | Please specify minimum detectable ΔA together with acquisition time, averaging conditions, repetition rate and confidence level. If you prefer, you may alternatively specify baseline noise at 1 s averaging and 10 kHz repetition rate.
32 | B | Please specify the chromatic focal shift across the probe wavelength range. (lower is better).
33 | A | Delay Line with following characteristics: · (i) max delay: > 5 ns, · (ii) resolution < 10 fs, · (iii) Power throughput: ≥ 80 %
34 | B | Please specify maximum delay in ns (higher is better)
35 | B | Please specify beam pointing drift at the sample over the full delay range (lower is better)
36 | B | Please specify max delay stage speed and acceleration (higher is better)
37 | B | Please specify whether there is automated alignment of the delay channel. (Systems with automatic alignment will be scored more highly. If alignment is not automated, please specify procedure. Simpler procedures are favoured).
38 | B | Please specify whether a reference channel is used. Please explain reason.
39 | B | Please describe delay-line optics and expected pulse broadening.
40 | A | UV/Vis detector with following characteristics: · (i) CMOS, NMOS or similar (please specify), · (ii) Pixels >= 256, · (iii) Resolution: ≤ 4 nm,
41 | B | Please state number of pixels in sensor
42 | B | Please state maximum acquisition rate with associated effective pixel count (higher rates and effective pixel counts favoured). If several pixels are binned together in the acquisition, this counts as one effective pixel.
43 | B | Please specify resolution of UV/Vis detector (lower is better)
44 | B | Please specify cooling method for UV/Vis detector
45 | B | Please specify dynamic range of UV/Vis detector (higher is better)
46 | B | Please specify read noise of UV/Vis detector (lower is better)
47 | B | Please specify acquisition rate of UV/Vis detector in scans per second (higher is better)
48 | A | · NIR detector with following characteristics: · (i) InGaAs or similar (please specify), · (ii) resolution: ≤ 15 nm (lower is better), · (iii) acquisition rate: > 3000 per second
49 | B | Please specify resolution of NIR detector (lower is better)
50 | B | Please specify cooling method for NIR detector
51 | B | Please specify dynamic range of NIR detector (higher is better)
52 | B | Please specify read noise of NIR detector (lower is better)
53 | B | Please specify acquisition rate of NIR detector in scans per second (higher is better)
54 | A | Long lifetime extension with the following characteristics: · (i) IRF: ≤1 ns (shorter better), · (ii) temporal overlap with fs measurement regime, · (iii) time-scales up to at least 400 microseconds, · (iv) pumping conditions consistent with ultrafast measurements.
55 | B | Please specify IRF of extension (shorter better)
56 | B | Please specify longest accessible timescale of extension (higher is better)
57 | B | Please specify spectral range of long lifetime extension. (wider range better)
58 | B | Confirm whether any manual alignment is needed to switch between ultrafast and long lifetime configurations. (Systems requiring minimal or no manual optical realignment will be scored more highly.).
59 | A | Control software with the following characteristics: · (i) Every individual kinetic scan saved, so previous scans are not lost if experiment is aborted · (ii) Two levels of user access for increased user-friendliness - basic (default, most commonly used settings) - advanced (allows to change DAQ parameters, such as continuum stability thresholds, etc.) · (iii) fully documented API (Application Programming Interface) with example programs. · (iv) API should be callable from MATLAB and Python.
60 | B | Please specify whether system offers support for multiple choppers to facilitate customized experiments, e.g.: - "pump-pump-probe” - “pump-dump-probe” or experiments where the probe beam is also modulated. (Capability preferred)
61 | B | State whether API can be called from any programming language (yes - preferred)
Computer/Software |  |  | Yes | Partly | No | Computer/sofware
62 | A | Critical Updates: All available updates relating to security and "bug fixes" should be delivered free of charge over the equipment's expected lifetime. Other software updates should be offered to the buyer based on the same premium model as the offer for this tender.
63 | B | Other Updates: Software updates should be offered to the buyer free of charge for the first two years after delivery. After that, they should be offered on the same premium model as the offer for this tender.
64 | B | PC: Please state whether a PC is included in the offer. If no PC is included, please state technical requirements for the PC.
65 | A/B | Licence: All delivered software to the requested equipment shall be included and specified. The buyer should have the right to perform installations/changes if necessary in the delivered equipment/systems. All required licences shall be included in the offer. If there are special conditions to the above, the supplier shall specify this. The buyer states also that we shall have the right to loan the procured equipment including the licenced software to a third party. If there are restrictions to such, these shall be specified.
66 | A | Intellectual property: If the supplier which the buyer has purchased right of use from, possesses the intellectual property in all delivered products, both hardware and software, special limitations to the buyers right to use shall be clearly described.
Environment |  |  | Yes | Partly | No | Environment
67 | B | The equipment should have a long service life. The supplier is asked to state the expected service life of the equipment and wear parts.
68 | B | The supplier is asked to account for the need for repairs, upgrades and spare parts for the equipment offered based on deliveries of similar equipment over the past 10 years.
69 | B | The equipment should have low energy consumption. The supplier is asked to account for and document electricity consumption.
70 | B | The supplier is asked to account for the equipment's "carbon footprint" from production to delivery. Documented calculations will be given priority.
71 | B | The supplier is asked to account for and document whether and which environmental labels and/or environmental standards the equipment meets.
|  | Options | Yes | Partly | No | Options
72 | B | The supplier shall also offer additional year(s) of service in appendix 3 (attached), which the buyer shall have the right, but not be obligated to, purchase. Please specify the conditions for extended service and quote price for both full service and maintance service in appendix 3, Table 3
73 | B | The supplier shall also offer additional year(s) of extended warranty in appendix 3 (attached), which the buyer shall have the right, but not be obligated, to purchase. Please specify the conditions for extended warranty and quote price in appendix 3, Table 3.
The buyer specifies that the supplier must elaborate/comment on all responses in the Response/Comments column. If the supplier considers it relevant, the supplier shall post attachments that underpin responses. Requirements which will be referred to as "refers to the attachment" and or incomplete response, may be perceived as not answered and could get 0 points in the evaluation.