2026028032 Scanning Electron Microscope (SEM)

Multipurpose, analytical low vacuum field emission SEM (scanning electron microscope) with EDS (Energy Dispersive X-ray Spectroscopy)

Part 2 – Annex B

Requirement specification

Introduction

The Norwegian Defence Laboratories, Chemistry and materials section, intends to enter an agreement regarding purchase of a Multipurpose, analytical low vacuum field emission SEM (scanning electron microscope) with EDS (Energy Dispersive X-ray Spectroscopy) as specified in Appendix 1.

Requirement table

The Purchaser's requirements

The Purchaser has listed their requirements in the following manner in Annex 1.

No.:	Unique serial number for the requirement

Description:	Description of the requirement

Documentation type:	Description of the documentation verifying the fulfilment of the requirement

Requirement type:	See the table below

| Absolute requirement: A | The requirement must be met. The requirement constitutes a minimum requirement which is not subject to relative assessment. |
| --- | --- |
| Absolute requirement: A* | The requirement must be met. The asterix indicates that the requirement is a minimum requirement, exceeding the requirement will be subject to evaluation under the relevant award criteria. |
| Requested requirement: B | The requirement should be met, but is not required. The requirement is subject to evaluation under the relevant award criteria. |

The Contractor's response

The Contractor shall respond to the Purchaser's requirements in the following manner in the tables in Annex 1:

Offered:	The extent to which the contractor is able to fulfil the requirement. Answered by yes, no or partially. If the requirement is responded to as partially fulfilled, the solution must describe the lack of fulfilment. If an A or an A* requirement is not fulfilled or partially fulfilled, this may entail that the contractor is rejected from the tender.

Solution:	Detailed information describing how the requirement is fulfilled.

If the Contractor e.g. due to lack of space, finds it impractical to fill in the solution in the requirement table, the description of the solution may be enclosed in a separate appendix. The requirement table shall indicate where the solution is and the solution shall clearly specify which requirement is detailed.

Technical requirements

| Req. | Priority: | Description: | The supplieres fulfilment: |  |  | Respons/comments: |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Yes | Partly | No |  |
| Delivery |  |  | Delivery |  |  |  |
| 1 | A | The offer shall include the following components, such as: - The delivery of all items included in the offer after formal order from the buyer - Documentation in English for the system/equipment - Software, which is to be the latest updated version at the time of delivery - List of consumable parts with prices. - Provision of all documents pertaining to the equipment which are part of this tender has to include, but not limited to, maintenance manuals, operation manuals and commissioning certificates. |  |  |  |  |
| 2 | A | Delivery: All costs related to the LVFESEM system up to Site Acceptancy Test (SAT) and formal takeover must be included in the offer. |  |  |  |  |
| 3 | A | Delivery dates for shipping, installation, Site Acceptancy Test (SAT) and training must be given. |  |  |  |  |
| 4 | A | The supplier must be able to present the offered solution for a demonstration if the buyer should choose to have the system demonstrated. The buyer should have the right to a demonstration but shall not be obligated to do so. The demonstration should either be presented at the buyers location and/or at a location where the offered system is being used. Please refer to sites where the offered system is being used/exhibited. |  |  |  |  |
| 5 | A | Service content and spare parts in the warranty for a period of 12 months from SAT must be described and included in the offer. |  |  |  |  |
| Imaging quality |  |  | Imaging quality |  |  |  |
| 6 | A | Resolution: Resolution at 15kV or higher must be 1 nm or better, resolution at 1kV must be 2.5nm or better, without using stage bias. Specify secondary electron image and backscatter electron image resolutions at 20 nA at analytical WD. Resolutions must be documented as part of the acceptance test. |  |  |  |  |
| Hardware |  |  | Hardware |  |  |  |
| 7 | A | The offer must contain a multipurpose Low Vacuum Field Emission Scanning Electron Microscope (LVFESEM) and detectors for standard SED, upper SED, LVSED, solid state BED (fully retractable), EDS (SDD) detector, chamber scope, specimen current meter/probe current detector and optical camera for specimen overview image and navigation. Specify type, model and locations of detectors included. The supplier can offer different combinations of LVFESEMs and detectors. Microscope must be compatible with analysis of magnetic samples. |  |  |  |  |
| 8 | A | Safety Devices: Protective devices against power failure must be included. In case the SEM electronics and/or lenses are water cooled, protection against water failure must be included. |  |  |  |  |
| 9 | A | Ergonomics: The instrument must have a combination of mouse control and buttons/knobs for the most common functions and trackball/stick for specimen movement, specify. |  |  |  |  |
| 10 | A | Electron gun: The electron source must be a Schottky field emission gun. FE-emitters, including expenses in connection with emitter exchange, for 10 years consumption must be included. |  |  |  |  |
| 11 | A | Electron gun: Specify the emitter system. Maximum probe current at 20 kV must be minimum 100 nA, specify objective aperture size at maximum probe current. |  |  |  |  |
| 12 | A | Scannable area: Minimum scannable imaging area in mm2 at 20kV and WD 20mm is 10 mm2 |  |  |  |  |
| 13 | A | Chamber and sample holders: The chamber must be compatible with the following detectors/equipment: SED, upper SED, solid state (retractable) BED, EDS detector, chamber video camera, NORDIF-EBSD detector*, Faraday cup, absorbed current meter, i.e. all commonly used equipment that can be added to the chamber. Stage navigation system must be included. The specimen stage must be eucentric. One extra set of standard specimen holders must be included, specify. * Must be transferred from current SEM during system installation. |  |  |  |  |
| 14 | A | Specimen stage and field of view: Chamber size, stage movement range and field of view are of very high importance, specify 14.1-14.5. Minimum requirements: - stage movement in x-y directions: 25cm2 - stage movement in z direction: 40mm - specimen height: 60mm - specimen weight: 1kg. - scannable imaging area at 20kV and WD 20mm: 10mm2 |  |  |  |  |
| 14.1 | B | Specify maximum stage movement in x-y directions in cm2. |  |  |  | 0; <25 cm2. 1; 26-50. 2; 51-75. 3; 76-100. 4; 101-125. 5; 126-150. 6; 151-175. 7; 176-200. 8; 201-225. 9; 226-250. 10; >250 cm2. |
| 14.2 | B | Specify maximum stage movement in z direction in mm. |  |  |  | 0; <40 mm. 1; 41-50. 2; 51-60. 3; 61-65. 4; 66-70. 5; 71-75. 6; 76-80. 7; 81-85. 8; 86-90. 9; 91-95. 10; >95 mm. |
| 14.3 | B | Specify maximum specimen height in mm. |  |  |  | 0; <60 mm. 1; 61-70. 2; 71-80. 3; 81-85. 4; 86-90. 5; 91-95. 6; 96-100. 7; 101-105. 8; 106-110. 9; 111-115. 10; >115. |
| 14.4 | B | Specify maximum specimen weight in kg. The sample size and weight should not severely affect the optimum performance of the instrument itself. |  |  |  | 0; <1 kg. 1; 1-1,40. 2; 1,41-1,80. 3; 1,81-2,20. 4; 2,21-2,60. 5; 2,61-3,00. 6; 3,01-3,40. 7; 3,40-3,80. 8; 3,80-4,20. 9; 4,21-4,60. 10; >4,60 kg. |
| 14.5 | B | Specify maximum scannable imaging area in mm2 at 20kV and WD 20mm. |  |  |  | 0; <10 mm2. 1; 11-14. 2; 15-18. 3; 19-22. 4; 23-26. 5; 27-30. 6; 31-34. 7; 35-38. 8; 39-42. 9; 43-46. 10; >46 mm2. |
| 15 | A | Chamber/stage/sample holders: Specify (by mechanical drawing) the chamber dimensions and location/geometry of ports for attachments. Specify the chamber detector compatibility. Specify ports, dimentions and positions by technical drawing. Specify standard sample holders sizes included. |  |  |  |  |
| 16 | A | EDS hard- and software to be included: - Peltier cooled EDS detector (SDD), minimum 60mm², specify resolution (MnKα). - Software: qualitative- and quantitative analysis. Point, line and area (mapping). Specify functions. - SEM and EDS systems must be fully integrated, specify. - System for automated qualitative and quantitative particle analysis, specify. - Minimum 5 off-line licenses for post processing must be included, specify. - MS Office on the EDS PC. |  |  |  |  |
| 17 | A | Low vacuum mode: Variable pressure up to minimum 100Pa. Specify influence on probe current and field of view. |  |  |  |  |
| 18 | A | EBSD detector: The FESEM must be compatible with the NORDIF EBSD pioneer detector since this detector will be transferred from current SEM. https://www.nordif.com/en/products/nordif-pioneer-uf-series |  |  |  |  |
| 19 | A | Accessories (specify equipment): - Air lock sample exchange system to avoid ventilating the main chamber. - Stand-alone plasma cleaner for sample preparation, specimen diameter >100mm, compatible with chamber and/or air lock mounting. - High resolution sputter coater (gold) for sample preparation. - Specimen holders for large/heavy specimens. |  |  |  |  |
| 20 | A | Display: 3 (minimum 27") high quality monitors must be included for SEM/EDS and software control. |  |  |  |  |
| 21 | A | Vacuum system: The vacuum system must include vacuum meters. The vacuum system must have a turbo-pump and come with a backup system as dry rotary/scroll pump(s) or a diaphragm pump(s). Specify the vacuum system. Pumping time (to get a SE image in high vacuum mode) after opening the front door (no specimen installed) must not exceed 5 minutes. |  |  |  |  |
| 22 | A | Accessories, spare parts and tools: One set of tools and spare parts necessary for user service and daily maintenance of the microscope and a set of standard samples for instrument calibration. Specify which accessories are included. |  |  |  |  |
| 23 | A | Electricity/power: The equipment is required to meet the Norwegian power supply system (240 V, 50 Hz). Describe the offered solution for electrical and other technical preferences. Requirements for UPS system to support vacuum system for 6h must be specified. |  |  |  |  |
| Software |  |  | Software |  |  |  |
| 24 | A | Specify the operating system software (OS) version for application PC(s). The supplier should have plans for upgrades of the instrument and/or the system offered, including the software and OS. Specify. |  |  |  |  |
| 25 | A | Image processing software: The system must have an image analysis software capable of image processing such as automatic stitching and 2D measurements. |  |  |  |  |
| 26 | A | Updates: All available updates which relates to security and "bugfix" must be delivered free of charge in the equipments' expected lifetime. |  |  |  |  |
| 27 | A | Licence: Licences for all the delivered softwares to the requested equipment shall be included and specified. The buyer must have the right to perform installations/changes if necessary in the delivered equipment/systems. All required licences shall be included in the offer. If there are special conditions to the above, the supplier shall specify this. |  |  |  |  |
| Training |  |  | Training |  |  |  |
| 28 | A | Training: Training for operation of the SEM shall be included in the offer. The offer shall include on-site training for 2 days for a minimum of 3 personnel from the buyer. The trainining must take place soon after SAT (Site Acceptance Test). In addition, both proffesional and technical staff from the buyer must have the option of being present at the time/date of installation and performance of the SAT (Site Acceptance Test). Additional 2 days of training for minimum 3 persons after 6-12 mnd after installation must be included. |  |  |  |  |
| Warranty |  |  | Warranty |  |  |  |
| 29 | A | Warranty: The system/equipment shall be offered with a minimum of 12 months of warranty. The warranty period shall begin after the Site Acceptancy Test (SAT). Content and coverage of the guarantee must be specified. Warranty period starts when the delivery has been accepted through a Site Acceptance Test (SAT). |  |  |  |  |
| Service |  |  | Service |  |  |  |
| 30 | A | Service and support: - All maintenance agreements must include telephone and online support for all types of services. If remote support is available for the offered equipment/system, this must also be included. - The supplier must dedicate named technicians with personal telephone number for contacting when needed. - A response time of no more than 72h on site is required. - The service contract must include one preventive maintenance visit and at least two emergency visits each year. - Travel expenses to Kjeller (Norway) must be included. - Spare parts are not to be included, except for 10 years consumption for the Schottky field emission gun FE-emitters., including expenses in connection with emitter exchange. |  |  |  |  |
| Spare parts |  |  | Spare parts |  |  |  |
| 31 | A | Service and support: Specify the number of years for spare parts delivery for the SEM and EDS systems after installation, minimum 10 years. |  |  |  |  |