## Requirements
Requirements Specification for procurement of X-Ray Fluorescense (XRF) Spectrometer
1. General Description 1.1 Purpose and Scope of the Procurement The purpose of this procurement is to acquire a benchtop X-ray Fluorescence (XRF) Spectrometer for elemental analysis of soil and food samples at NIBIO's research facilities. The instrument will be used for quantification of major and trace elements, including cadmium (Cd) and other environmentally relevant elements, to support research in soil quality assessment, contamination monitoring, nutrient analysis, and geochemical characterisation. The XRF system will form part of NIBIO's existing analytical sensor portfolio and will complement other measurement approaches, including spectroscopy and remote sensing. The instrument shall support the development of standardised methodologies for data acquisition, calibration, and multi-project data comparability. The equipment shall be suited for laboratory use at NIBIO Apelsvoll, Nylinna 226, 2849 Kapp, Norway. 1.2 Estimated Value The total estimated value of this procurement is approximately NOK 1,500,000 excl. VAT, including options. This is a non-binding estimate. The procurement includes non-exclusive options for the Buyer to acquire consumables, spare parts, and service agreements beyond the standard warranty period over the operational lifetime of the system. 1.3 Options The procurement includes non-exclusive options for: • Annual service and maintenance agreements beyond the warranty period • Additional training programmes and application support • Extended automation and high-throughput accessories • Consumables, replacement components, and calibration standards Option pricing shall be specified in the Price Schedule (Chapter 6 in this form).
2. Description of the Requirements Table The table below presents a comprehensive overview of both requirements specifications and evaluation criteria for the procurement of an XRF Spectrometer for NIBIO. Mandatory Requirements (M – green fill): These are the minimum specifications a supplier must satisfy for their offer to be considered. Mandatory requirements are non-negotiable. Any offer that does not meet all mandatory requirements will be excluded from further evaluation. Evaluation Criteria (E – yellow fill): These criteria are used to compare, and rank offers that meet all mandatory requirements. Evaluation criteria assess quality, performance, and value beyond the minimum threshold, and are weighted according to their relative importance in the final award decision. Supplier response instructions: Suppliers are required to complete the 'Supplier's Response' column for each row, either directly in the table or by providing a clear reference to a numbered attachment included with the offer. All attachments must be clearly labelled with the corresponding requirement ID. Incomplete responses may result in the offer not being evaluated.
Please state whether the requirement is fulfilled or not, and comment if necessary.
| Description | Type | Suppliers response (answer directly in the table or refer precicely to attached document)
1Technical Solution and Performance
1.1 | Instrument type and design The system shall be a benchtop, designed for high-throughput laboratory analysis of soil and food samples. The instrument shall be suitable for continuous laboratory operation. Confirm that the requirement is met. | M
1.2 | Analytical capability - quantitative analysis The system shall be capable of quantitative elemental analysis of major and trace elements in complex matrices, including soil and food samples. Confirm that the requirement is met. | M
1.3 | Radiation safety The instrument shall incorporate a fully radiation-protected enclosure in compliance with applicable safety standards and regulations for laboratory use in Norway. The system shall not require additional radiation shielding beyond the instrument enclosure under normal operating conditions. Confirm that the requirement is met. | M
1.4 | Sample type compatibility - prepared solid samples The system shall support analysis of prepared solid samples, including dried and ground powders, pressed pellets, and, where applicable, loose powders. Confirm that the requirement is met. | M
1.5 | High-throughput sample handling - automated loading The system shall support automated sample handling enabling continuous, high-throughput operation with minimal operator intervention. As a minimum, the instrument shall support the simultaneous loading and sequential analysis of at least 10 samples (e.g., via XY-stage loader, multi-position sample changer, or equivalent). Confirm that the requirement is met. | M
1.6 | Instrument control and data processing software A complete software package compatible with Windows 11 shall be included, providing full instrument control, sample positioning, spectra acquisition and evaluation, quantitative results output, and elemental composition visualization. All control software and hardware (PC or workstation) required for data processing shall be included in the offer. The software interface shall be in English. Confirm that the requirement is met. | M
1.7 | Electrical compatibility and CE marking The instrument shall be compatible with the Norwegian standard electrical system (230 V / 50 Hz, single phase) and shall be CE-marked at the time of delivery in accordance with all applicable EU/EEA directives. Confirm that the requirement is met. | M
1.8 | Limits of detection (LOD) - trace element performance Please provide a table of achievable LODs for detectable elements, under standard and optional measurement conditions. Preference will be given to instruments demonstrating the lowest achievable LODs for elements in soil and food matrices. | E
1.9 | Additional sample types and flexibility Please describe the full range of sample types that can be analysed with the offered system, including liquid samples, plant material, organic matter, and mixed/heterogeneous matrices, as well as any limitations or recommended preparation procedures. | E
1.10 | Sample holders and flexible sample dimensions Please describe the range of available sample holder types, cup sizes, and sample dimension limitations compatible with soil and food material formats. | E
1.11 | Please describe the automated sample handling solution, including maximum sample capacity and optional accessories. | E
1.12 | Calibration maintenance and drift correction Please describe the QC and drift correction capabilities of the system, including use of control reference samples and drift correction tools to maintain calibration stability across different measurement sessions and users. | E
1.13 | Instrument footprint, dimensions, and laboratory integration Please state the instrument dimensions (H x W x D) and weight, and describe footprint requirements including any ventilation, cooling, or dedicated infrastructure needed. | E
1.14 | Operational and consumable costs Please provide an estimate of expected monthly and annual operating costs (excluding staff time), including consumables, gases (if applicable), and routine maintenance items. Identify all consumable components requiring periodic replacement, including expected service lifetime and replacement frequency. | E
1.15 | Calibration strategies - standard and custom Please describe calibration strategies potentially available and suitable for different sample matrices (e.g., soil and plant-based materials). | E
1.16 | Data export formats The system is expected to support export of measurement results in standard, non-proprietary formats (e.g., CSV, XLSX, or equivalent). We expect the export of raw spectral data to also be supported. Please describe available data export formats and any limitations. | E
1.17 | Integrated safety features The system shall include all integrated safety interlocks required for safe XRF operation in a shared laboratory environment, including automatic X-ray shut-off upon opening of the sample enclosure. Please describe all built-in safety features. | E
2 Delivery, installation and training
2.1 | On-site installation Full on-site installation at NIBIO Apelsvoll (Nylinna 226, 2849 Kapp, Norway) by qualified manufacturer or authorized service personnel shall be included in the offer. The instrument shall be delivered functional and ready for use upon completion of installation. Confirm that the requirement is met. | M
2.2 | User training - included in offer Training in instrument operation and routine maintenance for a minimum of two (2) NIBIO personnel shall be included in the offer, covering instrument operation, sample preparation, software use, calibration procedures, and basic troubleshooting. Confirm that the requirement is met. | M
2.3 | Documentation - user manual and technical documentation All necessary documentation, including a complete user manual in English covering both operation and routine maintenance, shall be made available in electronic format upon completion of installation. Confirm that the requirement is met. | M
2.4 | Installation and training timeline Training shall take place no later than two (2) months after confirmed delivery and installation. Please describe the proposed plan and timeline for delivery, installation, and training. Earlier training is desirable | E
2.5 | Delivery date Please specify the expected delivery time from signed contract to delivery at place (DDP, Apelsvoll), expressed in weeks. Earlier delivery enabling earlier commencement of research activities will be rated favourably. | E
2.6 | Advanced and application-specific training options Please describe any available advanced or application-specific training programmes (e.g., calibration development for soil matrices, spectral analysis, method development) and associated costs. | E
3 Warranty, Service Level Agreement and customer support
3.1 | Warranty period - minimum requirement The warranty period shall be a minimum of twelve (12) months from the confirmed date of installation and acceptance, covering parts and labour for defects arising under normal use. Confirm that the requirement is met. | M
3.2 | Remote support - included in warranty The warranty period shall be a minimum of twelve (12) months from the confirmed date of installation and acceptance, covering parts and labour for defects arising under normal use. Confirm that the requirement is met. | M
3.3 | Application competence and technical guidance The supplier shall be able to guide NIBIO personnel in technical, software-related, and analytical questions related to soil and environmental XRF analysis, method development, and troubleshooting. Confirm that the requirement is met and describe the support structure. | M
3.4 | Extended warranty period A warranty period exceeding twelve (12) months is desirable but longer warranties are prefered. Please describe the length and scope of the warranty period offered and any applicable conditions. | E
3.5 | Annual service agreement - post-warranty Please describe the service level agreement the supplier can offer beyond the warranty period, including scope of services, response times, and conditions. Insert pricing in the Price Schedule (Chapter 6 of this form). | E
3.6 | Maintenance requirements and service infrastructure Please describe required maintenance and calibration routines, including a typical annual maintenance schedule. Describe the geographical availability and infrastructure of the supplier's technical service capability in Norway and the Nordic region. | E
3.7 | Additional value-adding features Solutions and features not explicitly requested but considered advantageous for NIBIO's intended use may be taken into account during evaluation. The supplier must document how and in what way these features provide benefit to the Buyer. | E
4 Environmental Award Criteria
4.1 | Energy efficiency Please provide documented energy consumption figures for normal operation and standby modes. Instruments demonstrating low energy consumption relative to analytical capability, or holding recognised certifications (e.g., Energy Star or equivalent), will be rated favourably. | E
4.2 | Instrument lifetime and durability Please provide manufacturer specifications regarding expected instrument lifetime, service intervals, and long-term performance. References from existing users or test data demonstrating durability will be viewed favourably. | E
4.3 | Environmental certifications Please document any relevant environmental certifications held by the instrument or manufacturer, including ISO 14001, RoHS, or equivalent recognised standards. Certifications demonstrating environmentally responsible manufacturing and operations of this instrument will be viewed favourably. | E
5 Options
5.1 | Extended range of the detector to measure light elements As an addition to the main procurement, we expect the device to have an option to extend the range of the measurable element to light elements starting from preferably carbon or nitrogen. If that option exist for the given device, please describe available/additional configurations for light element analysis. Where helium flushing or vacuum conditions are required, specify gas type, consumption rate, and estimated annual usage cost. Give price of the extention alone additional to the main procurement or if applicable give a total price of the extended setup that could be considered for purchase. Specify details. | E
6 Options: Price Schedule
6.1 | Annual Service and maintenance agreements beyond the warranty period
6.2 | Additional training programmes and application support
6.3 | Extended automation and high-throughput accesories
6.4 | Consumables, replacement components, and calibration standards