Table of Contents

1	Introduction	1

2	Background	2

3	Definitions	2

4	Compliance with requirements	2

5	Regulatory requirements	3

6	Delivery requirements	3

6.1	Delivery address for shipment purposes	3

7	Responsibilities	4

8	System description	4

8.1	General	4

8.2	System design	4

9	Requirements	4

9.1	General requirements	5

9.2	Functional and technical requirements	5

9.3	Regulatory health and safety requirements	6

9.4	Mechanical and electrical requirements	7

9.5	Service and maintenance requirements	7

9.6	Commissioning and training requirements	8

9.7	Deliverable data items	8

Introduction

This User Requirements Specification (URS) details the requirements for the purchase of a stationary instrument capable of dissolving refractory inorganic, as well as organic, matrices by acid digestion under high temperatures and pressure. The instrument is to be used as part of characterisation of radionuclides during safe operation and the awaiting decommissioning of the nuclear site in Halden, Norway. Thus, the instrument must be fit for purpose to ensure the safety of personnel and environment as well as satisfy the use of BAT.

The system will be installed at the laboratory at the following address:

NND – Norwegian Nuclear Decommissioning 
Dept. SFK - Radiation Protection, Physics & Chemistry

Tistedalsgata 20

1772 Halden, Norway

Background

The Department of Radiation Protection, Physics and Chemistry at NND, abbreviated NUKH SFK, is currently working on expanding the analytical capabilities of the laboratories to encompass characterisation of radionuclides in solid matrices. This requires specialised methods, due to the nature of refractory matrices and analytes requiring aggressive conditions to ensure complete dissolution with minimal chance for cross contamination prior to analysis.

Definitions

| NND |  | Norwegian Nuclear Decommissioning |
| --- | --- | --- |
| BAT | - | Best Available Technology |
| NUKH | - | Nuclear site at Halden |
| SFK |  | Dept. Radiation Protection, Physics & Chemistry |
| SLA |  | Service Level Agreement |
| EEA |  | European Economic Area |

Compliance with requirements

Potential Suppliers must indicate their compliance with all requirements in this URS by completing the appropriate column in the specification tables in section 9. In the case where a requirement is indicated as not fulfilled, the Supplier must describe how their offer deviates and, eventually, what the potential consequence is for the operation of the system.

The Supplier shall inform NND of any possible contradictory requirements in the URS so that these can be clarified and solved.

Any changes to or deviations from the requirements described in this URS must be formally approved by NND before they can be considered as valid.

The full specifications table shall be an appendix to the final contract.

Regulatory requirements

The instrument shall comply with all applicable Norwegian and EEA regulatory requirements for laboratory equipment and shall be CE marked. The Supplier shall provide the applicable EU Declaration of Conformity and a list of the EU/EEA directives, regulations and harmonised standards used for the CE marking and conformity assessment of the instrument.

The instrument shall comply with applicable product safety requirements, including requirements related to electrical safety, mechanical and structural safety, electromagnetic compatibility, pressure and high-temperature operation, and safe use in a laboratory environment involving radioactive sample handling. Materials and components exposed to chemicals shall be compatible with the intended use.

The Supplier shall provide sufficient documentation to demonstrate regulatory compliance, safe operation, and suitability for use in laboratories operating in accordance with ISO/IEC 17025, including safety information, user manuals, maintenance instructions, certificates, test documentation, and other relevant documentation specified in Section 9.7.

Delivery requirements

The Systems shall be delivered DDP according to Incoterms 2020 i.e., the Supplier is responsible for the transport and delivery of all parts of the system to the address specified in section 6.1, cleared for import with all taxes and duties paid.

The system shall be shipped with sufficient packaging to protect against weather, frost, theft and negligent handling. The Supplier is responsible for including transport insurance of the system.

The Supplier shall be present during the final installation of the systems, to provide guidance and/or collaborate with the End-User on all aspects of the final positioning of the different units of the system.

Should the Systems require special installation procedures that cannot be performed by our local engineers and electricians, it is expected that the Supplier provides the qualified personnel for the tasks.

Delivery address for shipment purposes

NND – Norwegian Nuclear Decommissioning

Attn: Camillo Bruhn

Dept. SFK - Radiation Protection, Physics & Chemistry

Tistedalsgata 20,

1772 Halden, Norway

Responsibilities

The Supplier is responsible for ensuring that the system is suitable for their intended use and shall not be exempted from this responsibility.

If the Supplier does not expressively identify any deviations from this specification, it is assumed that the delivery will be in full compliance with all requirements.

The Supplier is responsible for getting familiarised with the requirements laid out in relevant regulatory documents and standards that are referenced in this URS.

The offered systems shall be within the Supplier’s actual experience and normal scope of delivery.

System description

General

The system must be suitable for, and compatible with, analytical methods such as ICP-MS, LSC and alpha spectrometry for refractory solid materials typically found at and around nuclear reactor sites (e.g. concrete, corrosion-resistant alloys, and environmental samples such as river sediment and beach sand etc.)

The system shall not require connection to external networks, cloud services or vendor remote access for normal operation unless explicitly accepted by NND. Any network connection, remote access solution or data transfer interface shall be documented by the Supplier and approved by NND before use.

System design

The system is to be based on the technique commonly known as microwave assisted acid digestion using a single pressurised reaction chamber capable of withstanding pressures over 180 bar and 250 degrees Celsius. The system must be designed to withstand the routinely use of aqua regia and other commonly used mineral acids. The system is to be used to prepare samples prior to ultra-trace radionuclide analysis and thus must be designed in a way to minimise the chance for cross contamination during, and across, sample digestions.

Requirements

Must -requirements are mandatory. By answering "No" to such a requirement, the bid can be rejected.

Compliance with should -requirements will be a benefit in the competition.

Bidders are encouraged to answer briefly and precisely in the table below. References to relevant chapters, sections or pages in attachments may be included if necessary.

General requirements

| No. | Requirement specification | Must/ Should | Fullfill? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 9.1.1 | The system is to be placed on a laboratory bench and thus must be benchtop sized. The system size and wight cannot exceed the limits of a typical laboratory bench. | Must |  |  |
| 9.1.2 | Drawings of physical dimensions of the System’s detectors and instruments should be provided prior to delivery. | Should |  |  |
| 9.1.3 | Potential suppliers must state whether they can offer an optional service level agreement (SLA), as described in Section 9.5. | Must |  |  |

Functional and technical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 9.2.1 | The system must tolerate operational pressure and temperatures over 180 bar and 250o Celsius, respectively. | Must |  |  |
| 9.2.2 | The system must be designed to withstand the routinely use of aqua regia and other commonly used mineral acids. | Must |  |  |
| 9.2.3 | The system must be designed in a way to minimise the chance for cross contamination during – and a cross - sample digestions. Please specify. | Must |  |  |
| 9.2.4 | The instrument must have manual pressure release valve in case of power shortage. | Must |  |  |
| 9.2.5 | The exhaust port must be designed so that it can be connected to the laboratory ventilation system. All parts exposed to corrosive reaction gases must be made of corrosion-resistant materials. | Must |  |  |
| 9.2.6 | Effective and intuitive cleaning/maintenance procedures. Please specify. | Should |  |  |
| 9.2.7 | The pressure reaction chamber must have an analogue pressure gauge with values displayed in either metric units or atm. | Must |  |  |
| 9.2.8 | The System must be capable of normal operation without connection to the internet, cloud services or vendor remote access. | Must |  |  |
| 9.2.9 | The Supplier shall document all digital interfaces and communication functions of the System, including network ports, protocols, user access, data export functions and any remote access or cloud dependency. Any connection to NND’s laboratory intranet, external networks or other digital devices shall require prior approval by NND. | Must |  |  |
| 9.2.10 | The system must be delivered with all necessary instrument control and data acquisition software, including required licenses and interfaces, for configuration, operation, monitoring and export of data from the instrument. | Must |  |  |
| 9.2.11 | The system should operate at a noise level suitable for a laboratory environment. The supplier shall specify the noise level (dB(A)) during normal operation. Lower noise levels are preferred. | Should |  |  |

Regulatory health and safety requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 9.3.1 | All electrical equipment must be CE-marked and shall be delivered with compliance certificates. | Must |  |  |
| 9.3.2 | All safety warnings necessary for the safe operation of the system must be written in Norwegian or English. | Must |  |  |
| 9.3.3 | The instrument shall comply with applicable product safety requirements and safe use in a laboratory environment involving radioactive sample handling. | Must |  |  |

Mechanical and electrical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 9.4.1 | The System must be fully compatible with Norwegian power standards: 230 V supply voltage at 50 Hz. | Must |  |  |

Service and maintenance requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 9.5.1 | The warranty period must be at least 12 months from date of approved installation for the whole system. | Must |  |  |
| 9.5.2 | The Supplier should offer an annually renewable SLA, including corrective maintenance and on-site support during working days. The proposed SLA should specify response times, including acknowledgement of support requests within 24 hours during working days. | Should |  |  |
| 9.5.3 | Short response times for service are preferred. The Supplier should describe the response times available for service in Halden. Service personnel must be fluent in either Norwegian or English. Norwegian language skills are considered an advantage. | Should |  |  |
| 9.5.4 | The equipment offered must have an expected service life of at least 10 years. The supplier must guarantee support and the supply of spare parts for the device for at least 10 years after entering the contract. | Must |  |  |

Commissioning and training requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 9.6.1 | The Supplier should be present during final installation and commissioning of the System at the End-User location. | Should |  |  |
| 9.6.2 | The Supplier shall provide a comprehensive training in the set-up, daily use, and maintenance of the System at the End-User location. | Must |  |  |
| 9.6.3 | The Supplier should offer additional optional training courses beyond the initial commissioning training and describe the scope, duration, location and cost of such courses. | Should |  |  |
| 9.6.4 | The Supplier should offer application support for sample preparation of challenging matrices and describe the scope, availability and cost of such support. | Should |  |  |

Deliverable data items

| No. | Requirement specification | Must/ Should | Fulfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 9.7.1 | The System must be delivered with all necessary information required for its operation and routine maintenance. | Must |  |  |
| 9.7.2 | Written documentation must be delivered in English and/or Norwegian. | Must |  |  |
| 9.7.3 | The Supplier must provide documentation demonstrating that the instrument is suitable for use in laboratories operating in accordance with ISO/IEC 17025. The documentation shall include, where applicable, information on calibration procedures and intervals, preventive and routine maintenance, quality control measures, traceability of measurement data, and relevant validation or verification data supporting the intended use of the instrument. | Must |  |  |