Appendix 2 - Supplier’s solution description

# 1 Table of requirements

The requirements in the table of requirements form the basis for the Supplier’s description of its solution specification. The Supplier must give an extensive description of all requirements so that the Parties will have a common understanding of what is to be delivered, or which can give the Buyer alternative perspectives to the requirements expressed in the Buyer’s requirements specification.

The different requirements have different priority.

Priority A = 	Absolute requirements that must be fulfilled. Tenders that do not fulfil all the absolute requirements will be rejected.

Priority B = 	Important requirements that should be fulfilled, but it is not an absolute requirement. Answers will be of great importance for the evaluation of the tender. The requirements are weighted higher than C-requirements.

When filling out the table of requirements the Supplier shall mark an (X) in the column for yes, partly or no under “the Supplier’s fulfilment”. Reference to relevant documentation and any additional comments shall be made in the column “Comment”. It is reminded that it is the duty of the Supplier to document that it fulfils the set requirements. Any missing documentation will result in the requirement being considered as not fulfilled.

| No. | Priority | Description | The Supplier’s fulfilment |  |  | Comment |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Yes | Partly | No |  |
| 1 |  | Microscope and accessories |  |  |  |  |
| 1.1 | A | The offer includes a fully motorized and encoded research-grade inverted microscope. |  |  |  |  |
| 1.2 | A | The offer includes an anti-vibration table. |  |  |  |  |
| 1.3 | A | The microscope is equipped for eyepiece observation. |  |  |  |  |
| 1.4 | A | DIC is possible with at least one high-magnification lens. |  |  |  |  |
| 1.5 | A | The microscope is equipped with an LED light source for eyepiece observation, suitable for excitation of DAPI, EGFP, and mCherry. |  |  |  |  |
| 1.6 | A | DAPI, EGFP, and mCherry can be viewed individually or together by switching only the excitation wavelength. |  |  |  |  |
| 1.7 | A | The microscope is equipped with a motorized xy scanning stage, controllable via both joystick and software. |  |  |  |  |
| 1.8 | A | The stage has inserts to accommodate slides (75 x 25 mm), LabTek-type chambered coverglass, 35 mm and 60 mm Petri dishes, and multi-well plates (127 x 85 mm). |  |  |  |  |
| 1.9 | A | The microscope is equipped with a z-piezo stage. |  |  |  |  |
| 1.10 | A | The microscope is equipped with a focus stabilization mechanism to correct for focus drift. |  |  |  |  |
| 2 |  | Environment control |  |  |  |  |
| 2.1 | A | The microscope is equipped with incubation equipment to produce a humidified atmosphere of 5% CO₂ and ≤ 1% O₂ with all stage inserts provided. |  |  |  |  |
| 3 |  | Imaging |  |  |  |  |
| 3.1 | A | The microscope is equipped with a spinning disk confocal unit. |  |  |  |  |
| 3.2 | A | The spinning disk has pinhole patterns suitable for imaging of cell monolayers, 3D organoids, and small model organisms. |  |  |  |  |
| 3.3 | A | The following laser lines (or equivalent) are included: 405 nm, 488 nm, 555 nm, 640 nm. |  |  |  |  |
| 3.4 | A | A near infrared laser line (735 nm or equivalent) is included. |  |  |  |  |
| 3.5 | A | The system is equipped with all optical elements necessary to excite, detect, and reliably separate fluorochromes in blue, green, orange, red, and infrared. |  |  |  |  |
| 3.6 | A | The microscope can acquire and overlay brightfield images with fluorescence images. |  |  |  |  |
| 3.7 | A | The microscope is equipped with two high-end monochrome cameras. |  |  |  |  |
| 3.8 | A | The microscope is equipped with a color camera. |  |  |  |  |
| 4 |  | Objectives |  |  |  |  |
| 4.1 | A | A low-magnification lens (for overview images) is included. |  |  |  |  |
| 4.2 | A | A 10x plan apochromat lens is included. |  |  |  |  |
| 4.3 | A | A 20x plan apochromat lens is included. |  |  |  |  |
| 4.4 | A | A 20x lens suitable for thick plastic bottom plates is included. |  |  |  |  |
| 4.5 | A | A 20x plan-apochromat water immersion lens with large working distance is included. |  |  |  |  |
| 4.6 | A | A 40x plan apochromat water immersion lens with large working distance is included. |  |  |  |  |
| 4.7 | A | A 60x plan apochromat oil immersion objective is included. |  |  |  |  |
| 4.8 | A | A 100x plan apochromat oil immersion objective is included. |  |  |  |  |
| 5 |  | System PC and software |  |  |  |  |
| 5.1 | A | The system is equipped with a PC and monitor(s). |  |  |  |  |
| 5.2 | A | The PC has acquisition software installed to control the microscope, with functionality (modules and extensions) to perform imaging as described above (see section 2.1 “Intended use of the microscope and image analysis solution”) |  |  |  |  |
| 6 |  | Data storage and analysis solution |  |  |  |  |
| 6.1 | A | The system is delivered with a data storage and analysis solution, consisting of a powerful PC workstation or (preferably) a dedicated high-end storage and analysis solution with remote desktop functionality. |  |  |  |  |
| 6.2 | A | The data analysis solution provided includes software with functionality (modules and extensions) to perform image analysis as described above. (see section 2.1 “Intended use of the microscope and image analysis solution”) |  |  |  |  |
| 7 |  | Warranty, service, and training |  |  |  |  |
| 7.1 | A | The offer includes a maintenance visit and full system check and alignment immediately before (<1 month) the warranty period ends. |  |  |  |  |
| 7.2 | A | The offer includes a detailed plan for super user training on the microscope for at least two core facility staff members. |  |  |  |  |
| 7.3 | A | The offer includes a detailed plan for a course in image analysis for core facility users. |  |  |  |  |

# 2 Offered solution