Appendix 2 - Supplier’s solution description

# 1 Table of requirements

The requirements in the table of requirements form the basis for the Supplier’s description of its solution specification. The Supplier must give an extensive description of all requirements so that the Parties will have a common understanding of what is to be delivered, or which can give the Buyer alternative perspectives to the requirements expressed in the Buyer’s requirements specification.

The different requirements have different priority.

Priority A = 	Absolute requirements that must be fulfilled. Tenders that do not fulfil all the absolute requirements will be rejected.

Priority B = 	Important requirements that should be fulfilled, but it is not an absolute requirement. Answers will be of great importance for the evaluation of the tender. The requirements are weighted higher than C-requirements.

When filling out the table of requirements the Supplier shall mark an (X) in the column for yes, partly or no under “the Supplier’s fulfilment”. Reference to relevant documentation and any additional comments shall be made in the column “Comment”. It is reminded that it is the duty of the Supplier to document that it fulfils the set requirements. Any missing documentation will result in the requirement being considered as not fulfilled.

| No. | Priority | Description / Specifications | The Supplier’s fulfilment |  |  | Comment |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Yes | Partly | No |  |
|  |  | ABSOULTE REQUIREMENTS TABLE 1 |  |  |  |  |
| T1 - 1 | A | The maximal width of the table cannot exceed 171cm* |  |  |  |  |
| T1 – 2 | A | The maximal length of the table cannot exceed 100cm |  |  |  |  |
| T1 - 3 | A | Height can be adjustable, if not should be fixed in height between 80-100cm |  |  |  |  |
|  |  | ABSOULTE REQUIREMENTS TABLE 2 |  |  |  |  |
| T2 – 1 | A | The maximal width of the table cannot exceed 120cm* |  |  |  |  |
| T2 – 2 | A | The maximal length of the table cannot exceed 80cm |  |  |  |  |
| T2 - 3 | A | Height can be adjustable, if not should be fixed in height of approximately 90cm |  |  |  |  |
|  |  | ABSOULTE REQUIREMENTS TABLE 3 |  |  |  |  |
| T3 – 1 | A | The maximal width of the table cannot exceed 113cm* |  |  |  |  |
| T3 – 2 | A | The maximal length of the table cannot exceed 80cm |  |  |  |  |
| T3 - 3 | A | Height can be adjustable, if not should be fixed in height of approximately 90cm |  |  |  |  |

| No. | Priority | Description | The Supplier’s fulfilment |  |  | Comment |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Yes | Partly | No |  |
|  |  | SPECIFICATIONS ABSOULTE REQUIREMENTS |  |  |  |  |
| 1 | A | The table should support the use of a minimum of one (1) anesthesia gas induction chamber. |  |  |  |  |
| 2 | A | The induction chamber and mask(s) should be capable of operating both airflow and anesthesia concentration separately and simultaneously. |  |  |  |  |
| 3 | A | The table's surface and operating panels should be resistant to cleaning agents, including soap, Virkon S, and ethanol. |  |  |  |  |
| 4 | A | The table should function independently of the building’s ventilation system. |  |  |  |  |
| 5 | A | It should include a system to effectively remove excess anesthesia gas to prevent operator exposure (e.g., scavenger, gas evacuator system, or absorber). |  |  |  |  |
| 6 | A | If the system uses a carbon filter and pump, it should include a weight-based monitoring system with an alarm to indicate when the filter requires replacement. |  |  |  |  |
| 7 | A | Documentation should be provided to verify the system's effectiveness in proper gas evacuation. |  |  |  |  |
| 8 | A | The system should support the use of isoflurane anesthesia gas. |  |  |  |  |

| No. | Priority | Description | The Supplier’s fulfilment |  |  | Comment |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Yes | Partly | No |  |
|  |  | SPECIFICATIONS EVALUATION CRITERIA |  |  |  |  |
| 1 | B | Gas exposure: Describe how the system prevents and/or removes excess anesthesia gas exposure to the operator |  |  |  |  |
| 2 | B | Ease of Cleaning: Provide a detailed description of the design features that facilitate easy cleaning of the table. If the table includes a perforated surface, describe how it allows for cleaning underneath the surface. |  |  |  |  |
| 3 | B | Induction Chamber Design: Designed to minimize stress for mice, with appropriate color and shape to optimize animal welfare. Appropriately sized to accommodate adult mice weighing up to 60g, ensuring optimal animal welfare. - Describe the dimensions and design of the induction chamber, including: - The opening mechanism. - The color of the chamber. - The material used for construction. - Explain how the design optimizes usability and animal welfare. |  |  |  |  |
| 4 | B | Oxygen and Air Mixing System: - Provide a detailed description of how the table utilizes oxygen mixed with anesthesia gas. - Systems that can mix medical oxygen with isoflurane (e.g., via a compressor or pump) will be awarded the highest points. Clearly explain the mechanism used for mixing. - Indicate whether the system includes an option to use room air as an alternative to medical oxygen and describe how this functionality is implemented. |  |  |  |  |
| 5 |  | Anesthetic system: - Does the system support the use of both isoflurane and sevoflurane? If yes, please describe how it is possible to change between the two different gasses. - Describe volume of vaporizer and if there are different alternatives. - Describe how filling/refilling of anesthesia gas is done. - Describe if adjustment of flow for masks and induction chamber can be done separately. |  |  |  |  |
| 6 |  | Oxygen Flowmeter: - It is preferred that the table includes an oxygen flowmeter with adjustable gas flow settings (e.g., a display or flow meter for airflow adjustment) in the range of 0.1–2 L/min. - Describe the design and functionality of the flowmeter, if included. |  |  |  |  |
| 7 | B | Number of mice in anesthesia: - The table can accommodate a minimum of two (2) mice under surgical gas anesthesia using masks. - The anesthesia table preferably are capable of dosing gas anesthesia for up to three (3) mice simultaneously: one (1) for induction and two (2) for maintenance. |  |  |  |  |
| 8 | B | Anesthesia Masks: - Describe how the placement of masks can be adjusted on the table to accommodate different procedures and operators. - It is preferred that maintenance masks are fixed to the table during anesthesia to prevent movement to the animal during anesthesia. - Describe how this feature is implemented, if applicable. - Automatic mice recognition for adjustment of anesthesia gas (on/off). |  |  |  |  |
| 9 | B | Compatibility with Rats: - Describe whether the anesthesia system can be adapted or replaced with necessary equipment (e.g., induction chambers and masks) to ensure compatibility with rats. - Explain how the design, color, shape, and size of the equipment optimize animal welfare for rats. |  |  |  |  |
| 10 | B | Anesthesia Monitoring System: - Describe whether the table includes an option for anesthesia monitoring specifically for mice, including the following parameters/or indicate what available parameters: - Eg. ECG, pulse, respiration, temperature, SpO2, and expiratory CO2. - Describe if monitoring is possible for each mouse under mask anesthesia - Indicate whether the monitoring system is integrated into the table or available as an add-on. - It is preferred that the monitoring system includes adjustable alarm settings. - Describe the audible alarms if it possible to be muted or permanently disabled. |  |  |  |  |
| 11 | B | Heating System: The ideal table includes a heating system to maintain and adjust the body temperature of mice before, during and after anesthesia. Describe the heating system: - Can the heating be adjusted automatically using a feedback probe. - Does it include an override system for manual temperature adjustment. - Mechanism of action, type of heating, including avoidance of overheating. - Describe minimum and maximum temperature limits Or other designs if relevant. |  |  |  |  |
| 12 | B | Multi-Section Heated Work Surface: - Describe the possibility of a multi-section heated work surface, including: - At least one surface dedicated to procedures/surgery. - At least one surface dedicated to cages. |  |  |  |  |
| 13 | B | Lighting System: - The ideal table includes a lighting system with adjustable lux (brightness) and Kelvin (color temperature). - Indicate settings for lux and kelvin. Maximum and minimum levels. - It is a benefit if the light(s) are movable. Describe placement and design of the lights. - If relevant, describe how the lighting system provides optimal illumination for surgical procedures. |  |  |  |  |
| 14 | B | Microscope Compatibility: - One of the three (3) table ideally have room or is designed for the attachment or placement of an operating microscope. Describe what of the tables offer this solution and how it is implemented in the full setup. Describe whether the system supports image projection to a wall-mounted screen. |  |  |  |  |
| 15 | B | Mobility: - Describe whether the table is equipped with wheels, allowing it to be moved to a different lab if needed. |  |  |  |  |
| 16 | B | Gas and CO2 Removal: - Describe how the anesthesia system and tubing are designed to assist in the removal of excess anesthetic gases and expired CO2. - Describe how alarm/notification when filter is full. Provide details on the efficiency and safety of the gas removal system. |  |  |  |  |
| 17 | B | Euthanasia Chamber Compatibility: - Describe whether the system includes an option to connect a euthanasia chamber. The euthanasia chamber must allow for the gradual filling of CO2 in compliance with EU Directive 2010/63. |  |  |  |  |
| 18 | B | Separate preoperative and operative zones The table must be designed to clearly separate preoperative (unclean) zones from operative (clean/sterile) zones. |  |  |  |  |

# 2 Offered solution