Table of Contents

1	Introduction	2

2	Definitions	2

3	Compliance with requirements	3

4	Delivery requirements	3

4.1	Delivery address for shipment purposes	3

5	Responsibilities	4

6	System description	4

6.1	General	4

6.2	System design	4

6.3	Functional parameters	4

7	Requirements	5

7.1	Compliance Requirement	5

7.2	General requirements	5

7.3	Functional and technical requirements	6

7.4	Regulatory health and safety requirements	7

7.5	Mechanical and electrical requirements	8

7.6	Commissioning and training requirements	8

7.7	Deliverable data items	8

8	References	9

Introduction

This Requirements Specification details the requirements for purchase of a total of 20 (twenty) instruments for monitoring radioactive air particulates in indoor environments, each consisting of detector and signal processing instrument (reader instrument), including standard accessories and software for configuration and data acquisition (hereinafter referred to as “System”). Nineteen of these instruments will be stationary and one will be movable on a 2-wheel trolley.

The Systems will be used in different buildings for detection and measurement of indoor air particulates emitting α- and β- radiation in a shutdown nuclear research reactor awaiting decommissioning at the Institute for Norwegian Nuclear Decommissioning (NND) in Halden, Norway.

The Systems will be used at the following address:

NND – Norwegian Nuclear Decommissioning

Dept. SFK, Radiation Protection

Tistedalsgata 20

N-1772 Halden, Norway

This document covers requirements specifications related to:

- Air particulate detectors and their signal processing electronics (reader instruments).

- Dedicated software for configuration and data acquisition.

- Optional service level agreement (SLA) for urgent item replacements or repairs.

Definitions

CE	Conformité Européenne, Hallmark

NND	Norwegian Nuclear Decommissioning

NUKH SFK	Department of Radiation Protection, Chemistry and Physics, NND

URS	User Requirement Specifications

α	Alpha particle radiation

β	Beta particle radiation

N/A	Not Applicable

Compliance with requirements

Potential Suppliers must indicate their compliance with all requirements in this requirement specification, including completing the appropriate column in the specification tables in section 7.

Final tenders with substantial reservations to or deviations from the procurement documents, will be rejected from the competition.

Any deviations and/or reservations shall be specified in a precise and unequivocal manner in the tender, so that NND can assess them without contacting the tenderer. All deviations and reservations shall be made with a clear and unequivocal reference to the relevant point in the procurement documents and shall be stated in the tender form.

The tenderer shall state in a clear manner the consequences of any deviations and reservations for the performance of the contract, including its consequences for the price and/or other aspects of the tender.

The full specifications table in section 7 shall be an appendix to the final contract.

Delivery requirements

It is preferred that the Systems should be delivered DDP according to Incoterms 2020 i.e.. It is also preferred that the Supplier is responsible for the transport and delivery of all parts of the Systems to the address specified in section 4.1, cleared for import with all taxes and duties paid.

The Systems should be shipped with sufficient packaging to protect against weather, frost, theft and negligent handling. The Supplier should be responsible for including transport insurance of the Systems. 

Should the Systems require special installation procedures that cannot be performed by our local engineers and electricians, it is expected that the Supplier provides the qualified personnel for the tasks.

Delivery address for shipment purposes

NND – Norwegian Nuclear Decommissioning

Attn: Andre Remoe/Runar Ulseth

Dept. SFK, Radiation Protection

Tistedalsgata 20, N-1772 Halden, Norway

Responsibilities

The Supplier is responsible for ensuring that the Systems are suitable for their intended use and shall not be exempted from this responsibility.

If the Supplier does not expressively identify any deviations from this specification, it is assumed that the delivery will be in full compliance with all requirements.

The Supplier is responsible for getting familiarized with the requirements laid out in relevant regulatory documents and standards that are referenced in this document.

The offered Systems shall be within the Supplier’s actual experience and normal scope of delivery.

System description

General

The Systems must be suited to perform detection and measurement of α- and β-emitting radionuclides in air particulates with a high level of accuracy and sensitivity under continuous operation.

System design

The Systems are based on any radiation detection technology that fulfills the detection sensitivity and measurement requirements specified in this document. Detectors must be connected to their respective reader instruments for signal processing and local displaying of measurement values. The reader instruments must include analogue output ports of 4 mA – 20 mA that represent measured values. The analogue signals must be transmitted via cables up to 200 meters to our signal acquisition and processing system.

Functional parameters

The Systems must be capable of detecting a wide range of α- and β-radiation energy. A factory characterization of each of the detectors must be provided and ensure that the detectors’ efficiencies cover the stated energy range required.

Requirements

Bidders are encouraged to answer shortly and precisely in the table below and avoid referring to enclosures.

Must-requirements are minimum requirements. By answering "No" to such a requirement, the bid may be rejected after NND has made an assessment in accordance with the Norwegian procurement regulations.

Compliance with should-requirements will be a benefit in the competition.

Compliance Requirement

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.1.1 | The supplier shall describe compliance with suitable reference standards, such as IEC or equivalent standards, applicable to Category B equipment in accordance with IEC 61226 “Nuclear power plants – Instrumentation, control and electrical power systems important to safety – Categorization of functions and classification of systems” (Also known as NEK IEC 61226 in Norway). | Must |  |  |

General requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.2.1 | All cables necessary for the interconnection of different related units of each System must be included. | Must |  |  |
| 7.2.2 | Drawings of physical dimensions of the System’s detectors and instruments should be provided. | Should |  |  |
| 7.2.3 | The Systems will be placed in two separate buildings at the End-User location. The Supplier must be present during the final installation of the Systems, to provide guidance and/or collaborate with the End-User on all aspects of the final positioning of the different units of the System, including any lead shielding. | Must |  |  |

Functional and technical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.3.1 | 19 (nineteen) Stationary airborne particulate radioactivity monitors |  |  |  |
| 7.3.1.1 | Stationary instrument with integrated or separate detector for monitoring filter for air particulate contamination. Please specify max. cable length for offers with separate detector from instrument. | Must |  |  |
| 7.3.1.2 | The detector with particulate filter and instrument must be able to detect and measure α- and β-emitting particulates. | Must |  |  |
| 7.3.1.3 | The detector should be able to detect the following energy ranges. α-particulates: 3 MeV – 10 MeV. β -particulates: 50 keV – 2.5 MeV. | Should |  |  |
| 7.3.1.4 | The detector and instrument should cover at least the following measuring ranges. α-particulates: 0.01 – 10,000 Bq/m3. β-particulates: 1.0 – 1,000,000 Bq/m3. | Should |  |  |
| 7.3.1.5 | The particulate air filter should be based on moving filter technology, with periodic automatic filter changing configurable by the user. | Should |  |  |
| 7.3.1.6 | Air intake should be supplied through air hoses connected to the instrument/detector housing. Operational air flow rate should be between 10-60 liters/min. | Should |  |  |
| 7.3.1.7 | A flow meter for air intake must be included in the instrument, and the flow rate should be clearly visible. | Must |  |  |
| 7.3.1.8 | A dedicated air pump for air intake should preferably be part of the instrument. Please specify. | Should |  |  |
| 7.3.1.9 | The instrument must have at least one analogue output port of 4 mA – 20 mA for representing measurement values. | Must |  |  |
| 7.3.1.10 | The instrument’s analogue output of 4 mA – 20 mA should be able to deliver this current range via cable with shielded copper wires of 0.5 mm2 (20 AWG) up to 200 meters to a 250 ohms resistor, otherwise, a signal conditioner should be supplied. | Should |  |  |
| 7.3.1.11 | The instrument must be able to clearly display measurement values in Bq/m3 for α- and β-particulates detection. | Must |  |  |
| 7.3.1.12 | The instrument must be configurable for setting up alarm levels, i.e. H (high) and HH (high-high) for particulate detection, and alarm for high/low air flow rate. Please specify. | Must |  |  |
| 7.3.1.13 | The instrument must have a built-in visual and audible alarm. | Must |  |  |
| 7.3.1.14 | The instrument must include output for triggering preset alarms, i.e. relay output. Please specify. | Must |  |  |
| 7.3.1.15 | The detector and instrument should be easily calibrated by the End-User, and the requirement for calibration should not be more than once per year under normal operational conditions. | Should |  |  |
| 7.3.1.16 | The detector and instrument should function at temperatures from +5o to +40o Celsius. Please specify. | Should |  |  |
| 7.3.1.17 | The detector and instrument should have an ingress protection rating of at least IP 54 (dust protected and protected against splashing water). Please specify. | Should |  |  |
| 7.3.2 | 1 (one) mobile indoor air particulate monitor |  |  |  |
| 7.3.2.1 | Same requirements as specified under 7.2.1.2 to 7.2.1.17, except the cable requirement in 7.2.1.10. Instrument with or without detector must be mounted on a trolley that can be moved around and stationed at different locations. Please specify. | As in 7.3.1.2 to 7.3.1.17 |  |  |

Regulatory health and safety requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.4.1 | All electrical equipment must be CE-marked and shall be delivered with compliance certificates. | Must |  |  |
| 7.4.2 | All safety warnings necessary for the safe operation of the system must be written in English or in a Scandinavian language. | Must |  |  |

Mechanical and electrical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.5.1 | It is preferred that the reader instrument can be powered from a 24 Vdc power supply. | Should |  |  |
| 7.5.2 | If requirement 7.4.1 is not possible, the System should be fully compatible with Norwegian power standards: 230 V supply voltage at 50 Hz. | Must |  |  |

Commissioning and training requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.6.1 | The System, prior to shipment, should be tested for conformity with the manufacturer’s performance specifications and the minimum requirements specified herein. The results of the testing should be documented in a written protocol delivered together with the System. | Should |  |  |
| 7.6.2 | The Supplier must be present during final installation and commissioning of the System at the End-User location. This should include verifying key performance parameters as reported in the manufacturer’s specifications. | Must |  |  |
| 7.6.3 | The Supplier should provide a comprehensive training in the set-up, daily use, and maintenance of the System at the End-User location. | Should |  |  |

Deliverable data items

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.7.1 | The System must be delivered with all necessary information required for its operation and routine maintenance. | Must |  |  |
| 7.7.2 | Written documentation should be delivered in electronic format (i.e. pdf) in English and/or a Scandinavian language. | Should |  |  |

References

Reference 1:	EU Machinery directive (2006/42/EC)

Reference 2:	EU Low Voltage Directive (2014/35/EU)

Reference 3:	EU Electromagnetic Compatibility Directive (2014/30/EU)

Reference 4:	IEC 61226:2020 Nuclear power plants – Instrumentation, control and electrical power systems important to safety – Categorization of functions and classification of systems

Reference 5: 	ISO 18589-3:2015 Measurement of radioactivity in the environment — Soil — Part 3: 		Test method of gamma-emitting radionuclides using gamma-ray spectrometry

Reference 6:	ISO20042:2019 Measurement of radioactivity — Gamma-ray emitting radionuclides 		— Generic test method using gamma-ray spectrometry

Reference 7:	ISO11929 (part 1):2019 Determination of the characteristic limits (decision 			threshold, detection limit and limits of the coverage interval) for measurements of 		ionizing radiation — Fundamentals and application