Table of Contents

1	Introduction	2

2	Definitions	2

3	Compliance with requirements	3

4	Delivery requirements	3

4.1	Delivery address for shipment purposes	3

5	Responsibilities	4

6	System description	4

6.1	General	4

6.2	System design	4

6.3	Functional parameters	4

7	Requirements	5

7.1	Compliance Requirement	5

7.2	General requirements	5

7.3	Functional and technical requirements	6

7.4	Regulatory health and safety requirements	7

7.5	Mechanical and electrical requirements	7

7.6	Commissioning and training requirements	8

7.7	Deliverable data items	8

8	References	9

Introduction

This Requirements Specification details the requirements for the purchase of a total of two (2) stationary instruments for detection of radioactive Kr-85 gas in indoor environments. Both monitors consist of detector and measurement/reader instrument, including standard accessories and software for configuration and data acquisition (hereinafter referred to as “System”).

The Systems will be used in a shutdown nuclear research reactor awaiting decommissioning at the Norwegian Nuclear Decommissioning (NND) in Halden, Norway.

The Systems will be used by the department at the following address:

NND – Norwegian Nuclear Decommissioning 
Dept. SFK, Radiation Protection

Tistedalsgata 20

1772 Halden, Norway

The URS covers requirements related to:

- Kr-85 gas detectors and their signal processing electronics (reader instruments).

- Dedicated software for configuration and data acquisition.

- Optional service level agreement (SLA) for urgent item replacements or repairs.

Definitions

CE	Conformité Européenne, Hallmark

H-3	Tritium gas

NND	Norwegian Nuclear Decommissioning

Kr-85	Krypton-85 gas

NUKH SFK	Department of Radiation Protection, Chemistry and Physics, IFE

β	Beta (radiation)

ɣ	Gamma (radiation)

N/A	Not Applicable

Compliance with requirements

Potential Suppliers must indicate their compliance with all requirements in this requirement specification, including completing the appropriate column in the specification tables in section 7.

Final tenders with substantial reservations to or deviations from the procurement documents, will be rejected from the competition.

Any deviations and/or reservations shall be specified in a precise and unequivocal manner in the tender, so that NND can assess them without contacting the tenderer. All deviations and reservations shall be made with a clear and unequivocal reference to the relevant point in the procurement documents and shall be stated in the tender form.

The tenderer shall state in a clear manner the consequences of any deviations and reservations for the performance of the contract, including its consequences for the price and/or other aspects of the tender.

The full specification tables under section 7 shall be an appendix to the final contract.

Delivery requirements

It is preferred that the Systems should be delivered DDP according to Incoterms 2020 i.e. It is also preferred that the Supplier is responsible for the transport and delivery of all parts of the Systems to the address specified in section 4.1, cleared for import with all taxes and duties paid.

The Systems should be shipped with sufficient packaging to protect against weather, frost, theft and negligent handling. The Supplier should be responsible for including transport insurance of the Systems. 

Should the Systems require special installation procedures that cannot be performed by our local engineers and electricians, it is expected that the Supplier provides the qualified personnel for the tasks.

Delivery address for shipment purposes

NND – Norwegian Nuclear Decommissioning

Attn: Andre Remoe/Runar Ulseth

Dept. SFK, Radiation Protection

Tistedalsgata 20, N-1772 Halden, Norway

Responsibilities

The Supplier is responsible for ensuring that the Systems are suitable for their intended use and shall not be exempted from this responsibility.

If the Supplier does not expressively identify any deviations from this specification, it is assumed that the delivery will be in full compliance with all requirements.

The Supplier is responsible for getting familiarized with the requirements laid out in relevant regulatory documents and standards that are referenced in this URS.

The offered systems shall be within the Supplier’s actual experience and normal scope of delivery.

System description

General

The Systems must be suited to perform detection and measurement of Kr-85 radionuclide in air with a high level of accuracy and sensitivity under continuous operation.

System design

The Systems are based on any radiation detection technology that fulfills the detection sensitivity and measurement requirements. Detectors are either connected to or integrated with their respective reader instruments for signal processing and displaying measurement values. The reader instruments include standard process analogue output ports of 4 mA – 20 mA. The analogue signal, representing measurement values, is transmitted via cables of up to 200 meters to a centralized data acquisition and processing system.

Functional parameters

The Systems should be capable of detecting radiation energy emitted by Kr-85 nuclide. A factory characterization for each of the detectors shall be provided and ensure that the detectors’ efficiencies cover the stated energy range required.

Requirements

Bidders are encouraged to answer shortly and precisely in the table below and avoid referring to enclosures.

Must-requirements are minimum requirements. By answering "No" to such a requirement, the bid may be rejected after NND has made an assessment in accordance with the Norwegian procurement regulations.

Compliance with Should-requirements will be a benefit in the competition.

Compliance Requirement

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.1.1 | The supplier shall describe compliance with suitable reference standards, such as IEC or equivalent standards, applicable to Category B equipment in accordance with IEC 61226 “Nuclear power plants – Instrumentation, control and electrical power systems important to safety – Categorization of functions and classification of systems” (Also known as NEK IEC 61226 in Norway). | Must |  |  |

General requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.2.1 | All cables necessary for the interconnection of different related units of each System must be included. | Must |  |  |
| 7.2.2 | Drawings of physical dimensions of the System’s detectors and instruments should be provided. | Should |  |  |
| 7.2.3 | The Systems will be placed in two separate buildings at the End-User location. The Supplier must be present during the final installation of the Systems, to provide guidance and/or collaborate with the End-User on all aspects of the final positioning of the different units of the System, including any lead shielding. | Must |  |  |

Functional and technical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.3.1 | Two (2) stationary indoor Kr-85 gas monitors |  |  |  |
| 7.3.1.1 | Stationary instrument with integrated or separate detection chamber for high detection of Kr-85 gas contamination in air. Detection of other noble gases, especially H-3 and Rn-222, should be as low as possible. Please specify max. cable length for offers with separate detection chamber instrument from reader instrument. | Must |  |  |
| 7.3.1.2 | The detection chamber must include an air filter at the air intake for removing impurities. It must be easily replaceable by the End-User. | Must |  |  |
| 7.3.1.3 | The detector and reader instrument should cover a minimum measuring range from 1 kBq/m3 – 1 GBq/m3 with a max. deviation of ±10%. Please specify. | Should |  |  |
| 7.3.1.4 | The air intake to the detection chamber must be supplied through air sample hoses connected to it. Please specify operational flowrate in liters/min. | Must |  |  |
| 7.3.1.5 | A flow meter for air intake must be included in the reader instrument where the flow rate should be clearly visible. | Must |  |  |
| 7.3.1.6 | An integrated or separate air sampling pump should be part of the System. Please specify. | Should |  |  |
| 7.3.1.7 | The reader instrument must have at least one analogue signal output port of 4 mA – 20 mA for representing measurement values in logarithmic base 10 scale, configurable by the user. | Must |  |  |
| 7.3.1.8 | The reader instrument’s analogue output of 4 mA – 20 mA should be able to deliver this current range via cable with shielded copper wires of 0.5 mm2 (20 AWG) up to 200 meters and a 250 ohm resistor, otherwise, a signal conditioner should be supplied. | Should |  |  |
| 7.3.1.9 | The reader instrument must be able to clearly display measurement values in Bq/m3. | Must |  |  |
| 7.3.1.10 | The reader instrument must be configurable for setting up alarm levels, i.e. H (high) and HH (high-high) for Kr-85 detection and alarm for high/low air flow rate. Please specify. | Must |  |  |
| 7.3.1.11 | The reader instrument must have built-in visual and audible alarm. | Must |  |  |
| 7.3.1.12 | The reader instrument should include output for triggering preset alarms, i.e. relay outputs. Please specify. | Should |  |  |
| 7.3.1.13 | The System should tolerate operational temperatures from 5o to 40o Celsius. Please specify. | Should |  |  |
| 7.3.1.14 | The System should be easily calibrated by the user, and the requirement for calibration should not be more than once a year under normal operational conditions. | Should |  |  |
| 7.3.1.15 | The System should have an ingress protection rating of at least IP 54 (dust protected and protected against splashing water). Please specify. | Should |  |  |

Regulatory health and safety requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.4.1 | All electrical equipment must be CE-marked and shall be delivered with compliance certificates. | Must |  |  |
| 7.4.2 | All safety warnings necessary for the safe operation of the system must be written in English or in a Scandinavian language. | Must |  |  |

Mechanical and electrical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.5.1 | It is preferred that the reader instrument can be powered from a 24 Vdc power supply. | Should |  |  |
| 7.5.2 | If requirement 7.4.1 is not possible, the System must be fully compatible with Norwegian power grid standard: 230 V supply voltage at 50 Hz. | Must |  |  |

Commissioning and training requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.6.1 | The System, prior to shipment, should be tested for its conformity with the manufacturer’s performance specifications and the minimum requirements specified herein. The results of the testing should be documented in a written protocol delivered together with the System. | Should |  |  |
| 7.6.2 | The Supplier should be present during final installation and commissioning of the System at the End-User location. This should include verifying key performance parameters as reported in the manufacturer’s specifications. | Should |  |  |
| 7.6.3 | The Supplier should provide a comprehensive training in the set-up, daily use, and maintenance of the System at the End-User location. | Should |  |  |

Deliverable data items

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.7.1 | The System must be delivered with all necessary information required for its operation and routine maintenance. | Must |  |  |
| 7.7.2 | Written documentation should be delivered in electronic format (ex. pdf) in English and/or a Scandinavian language. | Should |  |  |

References

Reference 1:	EU Machinery directive (2006/42/EC)

Reference 2:	EU Low Voltage Directive (2014/35/EU)

Reference 3:	EU Electromagnetic Compatibility Directive (2014/30/EU)

Reference 4:	IEC 61226:2020 Nuclear power plants – Instrumentation, control and electrical power systems important to safety – Categorization of functions and classification of systems

Reference 5: 	ISO 18589-3:2015 Measurement of radioactivity in the environment — Soil — Part 3: 		Test method of gamma-emitting radionuclides using gamma-ray spectrometry

Reference 6:	ISO20042:2019 Measurement of radioactivity — Gamma-ray emitting radionuclides 		— Generic test method using gamma-ray spectrometry

Reference 7:	ISO11929 (part 1):2019 Determination of the characteristic limits (decision 			threshold, detection limit and limits of the coverage interval) for measurements of 		ionizing radiation — Fundamentals and application