Table of Contents

1	Introduction	2

2	Definitions	2

3	Compliance with requirements	3

4	Delivery requirements	3

4.1	Delivery address for shipment purposes	3

5	Responsibilities	4

6	System description	4

6.1	General	4

6.2	System design	4

6.3	Functional parameters	4

7	Requirements	5

7.1	Compliance Requirement	5

7.2	General requirements	5

7.3	Functional and technical requirements	6

7.4	Regulatory health and safety requirements	7

7.5	Mechanical and electrical requirements	8

7.6	Commissioning and training requirements	8

7.7	Deliverable data items	8

8	References	9

Introduction

This Requirements Specification details the requirements for purchase of a total of 10 (ten) Waterborne Particulate Radioactivity Monitors, each consisting of detector, detector pre-amplifier, detector high-voltage (HV) power supply (if needed), signal processing instrument (reader instrument), cabling and standard accessories and software for configuration and data acquisition (hereinafter referred to as “System”).

The Systems will be stationary and used for detection of direct ɣ-radiation from radioactive particulates in water flowing through installed water sample tanks. The Systems will be installed in a shutdown nuclear research reactor awaiting decommissioning at the Norwegian Nuclear Decommissioning (NND) in Halden, Norway.

The System will be used at the following address:

NND – Norwegian Nuclear Decommissioning

Dept. SFK, Radiation Protection

Tistedalsgata 20

1772 Halden, Norway

This URS covers requirements related to:

- NaI scintillation detectors and their signal processing electronics (reader instruments).

- Dedicated software for configuration and data acquisition.

- Optional service level agreement (SLA) for urgent item replacements or repairs.

Definitions

CE	Conformité Européenne, Hallmark

CPS	Counts per Second

HV	High voltage

N/A	Not applicable

NND	Norwegian Nuclear Decommissioning

NaI	Sodium iodide

NUKH SFK	Department of Radiation Protection, Chemistry and Physics, IFE

URS	User Requirement Specification

ɣ	Gamma (radiation)

Compliance with requirements

Potential Suppliers must indicate their compliance with all requirements in this requirement specification, including completing the appropriate column in the specification tables in section 7.

Final tenders with substantial reservations to or deviations from the procurement documents, will be rejected from the competition.

Any deviations and/or reservations shall be specified in a precise and unequivocal manner in the tender, so that NND can assess them without contacting the tenderer. All deviations and reservations shall be made with a clear and unequivocal reference to the relevant point in the procurement documents and shall be stated in the tender form.

The tenderer shall state in a clear manner the consequences of any deviations and reservations for the performance of the contract, including its consequences for the price and/or other aspects of the tender.

The full specification tables under section 7 shall be an appendix to the final contract.

Delivery requirements

It is preferred that the Systems should be delivered DDP according to Incoterms 2020 i.e.. It is also preferred that the Supplier is responsible for the transport and delivery of all parts of the Systems to the address specified in section 4.1, cleared for import with all taxes and duties paid.

The Systems should be shipped with sufficient packaging to protect against weather, frost, theft and negligent handling. The Supplier should be responsible for including transport insurance of the Systems. 

Should the Systems require special installation procedures that cannot be performed by our local engineers and electricians, it is expected that the Supplier provides the qualified personnel for the tasks.

Delivery address for shipment purposes

NND – Norwegian Nuclear Decommissioning

Attn: Andre Remoe/Runar Ulseth

Dept. SFK, Radiation Protection

Tistedalsgata 20, N-1772 Halden, Norway

Responsibilities

The Supplier is responsible for ensuring that the Systems are suitable for their intended use and shall not be exempted from this responsibility.

If the Supplier does not expressively identify any deviations from this specification, it is assumed that the delivery will be in full compliance with all requirements.

The Supplier is responsible for getting familiarized with the requirements laid out in relevant regulatory documents and standards that are referenced in this document.

The offered systems shall be within the Supplier’s actual experience and normal scope of delivery.

System description

General

The Systems must be suited to perform contamination detection and measurement of ɣ-ray emitting radionuclide particulates in water sample tanks with a high level of accuracy and sensitivity under continuous operation.

System design

The Systems are based on NaI scintillation detection technology that fulfills the detection sensitivity and measurement requirements. Detectors are installed inside existing water sample tanks in protective steel housing and are connected to their respective HV power supplies and reader instruments for signal processing and local displaying of measurement values. The connection between detectors and reader instruments is done with cables with lengths up to 15 meters without signal loss. The reader instruments include standard process signal output ports of 4 mA – 20 mA. The analogue signal, representing measurement values, is transmitted via cables of up to 200 meters to a centralized data acquisition and processing system.

Functional parameters

The Systems should be capable of detecting a wide range of ɣ-radiation energy. A factory characterization of each of the detectors should be provided and ensure that the detectors’ efficiencies cover the stated energy range required.

Requirements

Bidders are encouraged to answer shortly and precisely in the table below and avoid referring to enclosures.

Must-requirements are minimum requirements. By answering "No" to such a requirement, the bid may be rejected after NND has made an assessment in accordance with the Norwegian procurement regulations.

Compliance with Should-requirements will be a benefit in the competition.

Compliance Requirement

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.1.1 | The supplier shall describe compliance with suitable reference standards, such as IEC or equivalent standards, applicable to Category B equipment in accordance with IEC 61226 “Nuclear power plants – Instrumentation, control and electrical power systems important to safety – Categorization of functions and classification of systems” (Also known as NEK IEC 61226 in Norway). | Must |  |  |

General requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.2.1 | All cables necessary for the interconnection of different related units of each System must be included. | Must |  |  |
| 7.2.2 | Drawings of physical dimensions of the Systems’ detectors and instruments should be provided. | Should |  |  |
| 7.2.3 | The Systems will be placed in three separate buildings at the End-User location. The Supplier should be present during the final installation of the Systems, to provide guidance and/or collaborate with the End-User on all aspects of the final positioning of the different units of the System, including the lead shields. | Should |  |  |

Functional and technical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.3.1 | 10 (ten) stationary waterborne particulate radioactivity monitors |  |  |  |
| 7.3.1.1 | 6 (six) stationary 2” NaI scintillation detectors, including pre-amplifier, HV power supply for detector (if needed), and cable connection to a dedicated reader instrument as described in 7.2.1.8. | Must |  |  |
| 7.3.1.2 | The 2” NaI scintillation detector should have a sensitivity of around 1,600 CPS/µSv/h for Cs-137. Please specify. | Should |  |  |
| 7.3.1.3 | The 2” NaI scintillation detector must not be longer than 190 mm, and no larger than 61 mm in diameter. Please specify. | Must |  |  |
| 7.3.1.4 | 4 (four) stationary 3” NaI scintillation detectors, including pre-amplifier, HV power supply for detector (if needed), and cable connection to a dedicated reader instrument as described in 7.2.1.8. | Must |  |  |
| 7.3.1.5 | The 3” NaI scintillation detector should have a sensitivity of around 3,200 CPS/µSv/h for Cs-137. | Should |  |  |
| 7.3.1.6 | The 3” NaI scintillation detector must not be longer than 225 mm, and no larger than 88 mm in diameter. Please specify. | Must |  |  |
| 7.3.1.7 | The detectable radiation energy range by the NaI scintillation detectors must cover at least 100 keV – 2 MeV. Please specify. | Must |  |  |
| 7.3.1.8 | Ten stationary reader instruments that can be connected to the 2” or 3” NaI scintillation detectors specified. If the instruments exclude integrated HV power supply for the detectors, then separate HV power supplies must be included. Any configuration software must be included. | Must |  |  |
| 7.3.1.9 | The measurement range of the detectors with reader instruments should cover 1 CPS – 100,000 CPS. | Should |  |  |
| 7.3.1.10 | The reader instrument must include an analogue output port of 4 mA – 20 mA for representing measurement values in logarithmic base 10 scale, configurable by the user. | Must |  |  |

| 7.3.1.11 | The signal cable connecting the detector/pre-amp to the reading instrument and the cable connecting the detector to the HV power supply (if needed) should be up to 15 meters long without compromising signal quality. Specify optional cable lengths. | Should |  |  |
| --- | --- | --- | --- | --- |
| 7.3.1.12 | The reader instrument’s analogue output signal of 4 mA – 20 mA should be able to deliver this current range via cable with shielded copper wires of 0.5 mm2 (20 AWG) up to 200 meters and a 250 ohms resistor, otherwise, a signal conditioner should be supplied. | Should |  |  |
| 7.3.1.13 | The reader instrument must be configurable for setting up alarm levels, i.e. H (high) and HH (high-high). Please specify. | Must |  |  |
| 7.3.1.14 | The reader instrument must clearly display measurement values in CPS. | Must |  |  |
| 7.3.1.15 | The reader instrument should have a built-in visual and audible alarm. | Should |  |  |
| 7.3.1.16 | The reader instrument should include output for triggering preset alarms, i.e. relay output. Please specify. | Should |  |  |
| 7.3.1.17 | The detector must function at temperatures between 0o to +40o Celsius. Please specify. | Must |  |  |
| 7.3.1.18 | The reader instrument must function at temperatures between +5o to +40o Celsius. Please specify. | Must |  |  |
| 7.3.1.19 | The detector should have an ingress protection rating of at least IP 65 (dust tight and protected against water jets) or better. Please specify. | Should |  |  |
| 7.3.1.20 | The reader instrument with HV power supply (if applicable) should have an ingress protection rating of at least IP 54 (dust protected and protected against splashing water). | Should |  |  |

Regulatory health and safety requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.4.1 | All electrical equipment must be CE-marked and be delivered with compliance certificates. | Must |  |  |
| 7.4.2 | All safety warnings necessary for the safe operation of the system must be written in English or in a Scandinavian language. | Must |  |  |

Mechanical and electrical requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.5.1 | It is preferred that the reader instrument can be powered from a 24 Vdc power supply. | Should |  |  |
| 7.5.2 | If requirement 7.4.1 is not possible, the System should be fully compatible with Norwegian power standards: 230 V supply voltage at 50 Hz. | Should |  |  |

Commissioning and training requirements

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.6.1 | The System, prior to shipment, should be tested for its conformity with the manufacturer’s performance specifications and the minimum requirements specified herein. The test results should be documented in a written protocol delivered together with the System. | Should |  |  |
| 7.6.2 | The Supplier should be present during final installation and commissioning of the System at the End-User location. This should include verifying key performance parameters as reported in the manufacturer’s specifications. | Should |  |  |
| 7.6.3 | The Supplier should provide a comprehensive training in the set-up, daily use, and maintenance of the System at the End-User location. | Should |  |  |

Deliverable data items

| No. | Requirement specification | Must/ Should | Fullfills? Yes/No | Comment / response |
| --- | --- | --- | --- | --- |
| 7.7.1 | The System must be delivered with all necessary information required for its operation and routine maintenance. | Must |  |  |
| 7.7.2 | Written documentation should be delivered in electronic format (i.e. pdf) in English and/or a Scandinavian language. | Should |  |  |

References

Reference 1:	EU Machinery directive (2006/42/EC)

Reference 2:	EU Low Voltage Directive (2014/35/EU)

Reference 3:	EU Electromagnetic Compatibility Directive (2014/30/EU)

Reference 4:	IEC 61226:2020 Nuclear power plants – Instrumentation, control and electrical power systems important to safety – Categorization of functions and classification of systems

Reference 5: 	ISO 18589-3:2015 Measurement of radioactivity in the environment — Soil — Part 3: 		Test method of gamma-emitting radionuclides using gamma-ray spectrometry

Reference 6:	ISO20042:2019 Measurement of radioactivity — Gamma-ray emitting radionuclides 		— Generic test method using gamma-ray spectrometry

Reference 7:	ISO11929 (part 1):2019 Determination of the characteristic limits (decision 			threshold, detection limit and limits of the coverage interval) for measurements of 		ionizing radiation — Fundamentals and application